## Slide 1: Replacing the GD-70D Flow Sensor

## Slide 2: Replacing the Flow Sensor
Turn off power and remove the GD 70D from the mounting bracket.

## Slide 3: Replacing the Flow Sensor
Turn the GD-70D upside down to gain access of the flow sensor.
Replacement flow sensor is: P/N 31-1011-01.

## Slide 4: Replacing the Flow Sensor
Using a small Phillips screwdriver, remove the one screw that secures the flow sensor to the GD-70D.
Lift up on the connecting cable to remove sensor.

## Slide 5: Replacing the Flow Sensor
Lift out flow sensor and remove electrical cable from flow sensor.
Make sure that the old O-rings are removed from the chamber.

## Slide 6: Replacing the Flow Sensor
Old flow sensor removed.

## Slide 7: Replacing the Flow Sensor
Place new O-rings on the new flow sensor if not already installed.
Attach electrical cable to flow sensor and place flow sensor back into GD-70D.

## Slide 8: Replacing the Flow Sensor
Screw new flow sensor back into GD-70D.

## Slide 9: Replacing the Flow Sensor
Reattach the GD-70D to the mounting plate.
Turn on the GD-70D
Press and hold the MODE button.
First menu is 1-1, ZERO.
Press the UP button to Menu 1-3 Flow

## Slide 10: Replacing the Flow Sensor
With Menu 1-3 displayed, press the SET button to enter this mode.

## Slide 11: Replacing the Flow Sensor
In Menu 1-3 you will see the flow rate in Liters per Minute.
Press the Mode button to return to menu 1-3.

## Slide 12: Replacing the Flow Sensor
If flow rate is not at .5 LPM, use UP button to scroll to menu 1-8.
Push and release the SET button.

## Slide 13: Replacing the Flow Sensor
You will see dashes.
Press and HOLD the SET button to enter into Level 2 menu.

## Slide 14: Replacing the Flow Sensor
Using the UP button, scroll to Menu 2-5 DEF FLOW and press the SET button.

## Slide 15: Replacing the Flow Sensor
Push the SET button again to default flow.
This will bring you back to Menu 2-5.
Push and HOLD the MODE button to return to normal mode.

## Slide 16: Replacing the Flow Sensor
With instrument back in normal operation verify that flow indicator shows normal flow rate.
Calibrate the GD-70D when completed.