## Slide 1
MPS Flammable Sensor Training
Agenda:
What is Molecular Property Spectrometer (MPS)
Performance Vs. traditional sensors
TrueLEL
MPS overcoming short falls
How to test MPS
MPS Enablers
MPS Boundaries
Reliability testing
RKI Products MPS is adapted in

## Slide 2
Chemical resistant Ultem PEI housing
with integrated dust screen
Sensor Board (MPS + environ sensor)
Potting compound
CPU Board (MCU + drive electronics)
Interface board – 5 pin connectors
What Is MPS
The MEMS transducer is mounted on a PCB within a rugged enclosure open to ambient air.
Measurements of the thermodynamic properties of the combustible gas(s) mixture
Multiple transducer responses w/ environmental information is used to accurately determine the gas present
Provides accurate readings for multiple gases simultaneously that are linear from 0-100% LEL
Performs in condensing humidity conditions commonly encountered across the world

## Slide 3: MPS Versus Competing Technologies – Reported vs. Delivered
Drift & poisoning
(all curves)
CAT BEAD
(Based on typical datasheet k-factors)
Alarm set at 20%LEL
KEY TO THESE PLOTS
Typical NDIR
(Based on typical datasheet k-factors)
Over reports to gases not calibrated to: False Positive/Alarms
Blind to Hydrogen
Prone Error in Transient Temp/ Humidity Changes
Under reports to gases not calibrated to: Alarm to late/unsafe
Requires Frequent Servicing: Bump, Calibration, & Replacement
Short Lifetime - Poisoning
Higher false negative rate
Gas Delivered to Gas Reported
3
Notes: Almost every sensors data sheet has a similar plot of gas delivered to gas reported. You want to slope of 1:1
Below are same actual plots from our competitors.
NDIRs when calibrated to methane, are overly sensitive to other gases creating false positives, they are also blind to hydrogen, and are notoriously prone to error rate in rapid temp/humidity transients – timeout.
CatBeads are the opposite the tend to under report other gases – dangerous and their performance decays rapidly and requires constant maintenance and often times full replacement.

## Slide 4: MPS Versus Competing Technologies – Reported vs. Delivered
Alarm at 20%LEL
KEY TO THESE PLOTS
MPSTM Delivering TrueLELTM
MPS
Based on 3rd Party Testing and Empirical MPS Data
TrueLELTM: Accurate from 0-100% LEL Across all Common Combustible Gases, Including Hydrogen
Zero Maintenance: No Drift/Decay = No Calibration, No Replacement, No Correlation Factors
4

## Slide 5: The Data
Accuracy
  Top left shows our accuracy of the TrueLEL gases
  Top right shows traditional sensors
Traditional sensors characteristics
  Poisoning (Pellistor)
  Only accurate to one gas (Pellistor & NDIR)
  Cannot detect Hydrogen & Acetylene (NDIR)
  Humidity issues & false alarms (NDIR)
  Need regular calibration (Pellistor & NDIR)
  Narrower environmental window
Competitive matrix
  The bottom image shows NevadaNano matrix view
Complementary to your current portfolio
  Application
    Boiler rooms, Hydrogen economy, Water industry, Steel sector and Chemical sites
  Environmental
    -40 to 75 deg C upto 100% RH)
  Life span & total cost of ownership
    10+ years of life with no risk of poisoning and the sensor only needs to see bump testing gas to meet any SOP requirement

## Slide 6
MPS Technology Is Inherently Resistant To Poisons & Saturation
Typical Poisons
Silicones
Hydrogen Sulfide (H2S)
Ketones
Esters
Sulfur Containing Compounds
Higher Hydrocarbons
Typical Applications
Wastewater (H2S)
Workers using sealants, WD40, Armorial
Gas Utilities
Sour gas
Pellistor sensors can loose 95% accuracy response to methane, as a result their designs are not resistant to poisons
MPS will not be harmed if exposed to gas levels >100%LEL: when gas levels <= 100%LEL automatically/instantly reports concentration accurately. No recovery time

## Slide 7
MPS Technology Is Hydrogen Ready
The focus to find clean energy alternative from natural gas the usage of Hydrogen is rapidly increasing.
MPS Over traditional Cat-bead
Better accuracy
No poisoning
No saturation
No calibration or correction factor
Long life

## Slide 8
How To Test MPS Sensor

## Slide 9: RELIABILITY TESTING RESULTS & COMPLIANCE
 | Test | Specification | Conditions | Result
Hotplate | Accelerated Life: High Plate Temp | NNTS | 1.6/1.4/1.3X operation temperature | PASS: >40 years life expected
 | Accelerated Life: Fast Pulsing | NNTS | >118 million sensor temperature cycles | PASS: >7 years equiv. life demonstrated
Mech. | Drop | IEC 60068-2-31 | 1 meter | PASS
 | Shock | IEC 60068-2-27 | 50g @ 11ms | PASS
 | Vibration | IEC 60068-2-6 | 20g, 10Hz-2000Hz | PASS
Electrical | Electromagnetic Compatibility | CISPR 11 | Class B, Group 1 | PASS
 | Electromagnetic Immunity | IEC/EN 61000-4-3/-4/-5/-6/-8/-11 | See Spec. | PASS
 | Electrostatic Discharge | IEC/EN 61000-4-2 | See Spec. | PASS
 | Power Cycling | 12,000 cycles | 2 seconds on, 5 seconds off | PASS
Environmental | High Temp Operating | IEC 60068-2-2 | 1000 Hours @ 85C (Ambient RH) | PASS
 | Cold Temp Operating | IEC 60068-2-1 | 1000 Hours @ -50C (Ambient RH) | PASS
 | Heat/Humidity | IEC 60068-2-78 | Long duration @ 40C/93% RH | 
 | Sand/Dust | MIL-STD-810G | 510.5 | PASS
 | Ingress Protection | ATEX-default IP2X |  | PASS
 | Poisoning | NNTS | 1200 ppm-hours H2S (50 ppm for 24 hours) 100 ppm-hours HMDSO (10 ppm for 10 hours) 0.25 ppm-hours NO2 (3 ppm for 5 minutes) 0.83 ppm-hours HCN (10 ppm for 5 minutes) 0.75 ppm-hours SO2 (9 ppm for 5 minutes) 0.17 ppm-hours Cl2 (2 ppm for 5 minutes) 4.17 ppm-hours NH3 (50 ppm for 5 minutes) | PASS
Compliance | ATEX/IECEX Ex Protection | IEC 60079-0 | General Requirements for Ex Protection | PASS
 | ATEX/IECEX Ex Protection | IEC 60079-11 | Requirements for intrinsically safe 'i' protection | PASS
 | US Ex Protection | FM 3600/3610 | Similar to -0/-11 respectively but satisfies US requirements | PASS
 | Canada Ex Protection | CSA 22.2 60079-0/60079-11 | Similar to -0/-11 respectively but satisfies CA requirements | PASS
 | RoHS | 2011/65/EU & 2015/863 | Also includes China RoHS - SJT/T 11363 & 11364 | PASS
 | REACH | EC 1907/2006 (33 & 67) | See Spec. | PASS
Pass
Notes: Laundry list of industry standard tests. On the right is a plot of the raw transducer at highly increased pulse rate. Every 2 seconds to 10 times every second and we do it way above our normal operating temperature ~ 1000c
Equivalent of 10+ years

## Slide 10
MPS Boundaries

## Slide 11
MPS Flammable Sensor Enables:
No Field Calibration: MPS sensor technology does not require calibration or scheduled maintenance over it’s 10+ year life. This reduces interruptions to operations, with increased up-time. The comprehensive sequence of self-checks MPS automatically reports any problems with its operation. Along with reducing total cost of ownership associated to costly calibrations.
TrueLEL Multi Gas Accuracy: Numerous industries have multiple gas’s present. Traditional sensor technology can only accurately detect a single gas that it has been calibrated to. This can result in inaccurate readings, with false alarms, or even worse of no alarm. MPS sensor accurately detects 14 different combustible and automatically without the need for calibration or correction factors. The most accurate combustible gas sensor available today.
Sensor Poisoning Is a Thing of The past: Traditional sensors can be under the risk of poisoning in certain environments. Frequent testing is the only way to ensure performance is not being degraded. MPS sensor technology is designed for today's world, common poisons and contamination is resolved with MPS sensor technology.
Long Sensor life: 10+ year life expectancy reduces total cost of ownership by reducing cost associated with sensor replacement
Built In Environmental Sensor: Built-in temperature, humidity, and pressure sensor data are all calculated in real time. Zero false positives / drift.
Hydrogen Ready: Unlike NDIR technology

## Slide 12
Product Adaptation For MPS
Direct connect sensor to any of Beacon 110, Beacon 200, or Beacon 410A controllers
S2 Series
M2A Series