## Slide 1
GX-3R User Training

## Slide 2
Turn on the GX-3R by pressing and holding the POWER MODE button until you hear one beep
Turn on

## Slide 3
Press the POWER/MODE button to perform calibration.
Cal Reminder
Acknowledge and use
Perform a calibration
Press AIR button to continue without performing a calibration
If your instrument is past its 30 day calibration limit, you will see a CAL-LMT reminder at start-up

## Slide 4
After warm-up, the normal screen will appear.
Normal Mode

## Slide 5
Fresh Air
To perform a fresh air adjustment, make sure you’re in a fresh air environment.

## Slide 6
Press and hold the AIR button
Fresh Air

## Slide 7
When prompted, release the AIR button
Fresh Air
The GX-3R will beep once and return to normal mode after the fresh air adjust is complete

## Slide 8
To reset the instrument, go to a safe location. Once the readings return to fresh air values, then press the POWER/MODE button.
Alarm
In the event of an alarm, the GX-3R’s visual, audible and vibration alarms will alert you.
The readout will identify which sensor is in alarm.
The visual, audible, and vibration alarms will stop.

## Slide 9: Display Mode
To enter display mode, press and release the Power Mode button

## Slide 10
You will know you are in display mode if you see a display mode menu title on the bottom of your screen
Display mode vs. Normal Operation
PEAK
STEL
TWA
HC GAS LIST
USER ID
STN ID
CAL DATA
BUMP DATA
TEMPERATURE
ALARM-PT
To return to normal operation, scroll all the way through the display menu using the AIR button, or wait 20 seconds and the GX-3R will automatically return to normal mode

## Slide 11
Charging the instrument
Plug the charger into the instrument. A full charge takes 3 hours, and will power the instrument up to 25 hours.
A 30 minute charge will provide 8 hours of operation time.
Once the GX-3R is charged, the light will go from orange to green
Multi-unit AC Adapter is available (#49-0134)

## Slide 12: Calibration

## Slide 13: Calibration
If your instrument is past its 90 day (factory set default) calibration limit, you will see a CAL-LMT reminder at start-up
Press the AIR button to continue without performing a calibration
Acknowledge and use
Press the POWER/MODE button to perform calibration.
Perform a calibration

## Slide 14: Calibration
To perform a fresh air calibration
⎯ Make sure you are in a gas free environment
⎯ Press and hold the “AIR” button
⎯ When prompted, release the button
Note: If you are unable to perform a Fresh Air adjustment, it may be that “DEM ZERO” in Maintenance Mode has been turned off. In this case, you will have to enter User Mode and select GAS CAL  AIR CAL to perform one.

## Slide 15
Calibration
To perform a calibration without being prompted by the calibration reminder (CAL RMDR) turn the instrument off and restart it in User Mode
To enter User Mode:
Turn the GX-3R off
Press and hold the “AIR” button
Then press and hold the “POWER/MODE” button
Release after one beep

## Slide 16: Calibration
In the User Mode menu
  Use the “AIR” button to scroll options
  Use the “POWER/MODE” button to select
Select “GAS CAL”
  Note your sub-menu options are “AIR CAL” and “AUTO CAL”
  (NOTE: This is where you can do a “fresh air” calibration if you did not on start-up)
  Select “AUTO CAL”
  Calibration values will be displayed
  (LEL – 50%, O2 – 12%, CO – 50ppm, H2S – 25ppm) (factory set default values)

## Slide 17: Calibration
If the calibration values displayed do NOT match the gas cylinder values:
Enter the AUTO CAL menu (User Mode Gas Cal Auto Cal)
Scroll with “AIR” and select “CAL-P” with “POWER/MODE”
You should see the active channel marked with 3 dashes (- - -)
Use “AIR” to select which gas you want to change the value of
Press “POWER/MODE” to select
Use ”AIR” to adjust to the value matching your cylinder
(HINT: if you need to reverse the direction of your numbers, press AIR and POWER/MODE together quickly)
Press the “POWER/MODE” button to save
Press “AIR” to move to the next gas. Repeat as necessary
Select “ESCAPE” to return to the Auto Cal menu

## Slide 18: Calibration
Applying gas to the GX-3R
  Attach a fixed flow regulator to the calibration gas cylinder
  Attach the calibration adapter to the instrument and connect it to the 0.25 LPM regulator with tubing
  With “AUTO CAL” highlighted, press ”POWER/MODE”
  Turn on the regulator

## Slide 19: Calibration
  “AUTO CAL/CYL A” should be alternating on the display
  Press “POWER/MODE”
  Apply gas until readings stabilize or for 2 minutes maximum
  Press “POWER/MODE” Turn the gas regulator off

## Slide 20: Calibration
  If the calibration is successful, the unit will read “PASS”
  If the calibration is unsuccessful, the unit will read “FAIL”
Note that the fail message gives status for each sensor

## Slide 21: Calibration
AUTO CAL pass
  Remove the calibration adapter
  The unit will read “PASS,” then go through its start up sequence and enter measurement mode, ready for use
AUTO CAL fail
  The unit will read “FAIL” and return to the AUTO CAL menu in User Mode
  NOTE: In the “CAL EXPD” User Mode menu item, you can set the instrument to “CANT USE” to prevent the end user from continuing without calibration

## Slide 22
RP-3R Attachable Pump

## Slide 23
RP-3R Pump
Place the GX-3R right bottom first (as illustrated above) and then seat the left side until the clip latches.

## Slide 24
RP-3R Pump
The attachable probe has a dust filter to ensure the quality of your samples
The inlet filter housing contains a hydrophobic filter

## Slide 25
Turn off the GX-3R by pressing and holding the POWER MODE button
Turn off

## Slide 26
SDM-3R Calibration Station

## Slide 27
SDM-3R: Gas Ports
Air port:
Dust Filter
Gas 1:
4 gas mix
Gas 2:
Not used
Gas 3
Not used

## Slide 28
SDM-3R: USB port
USB port

## Slide 29
SDM-3R: Rear ports
AC power
Ethernet
PC Connection

## Slide 30
Connect Cylinder to SDM-3R
Connect a 4 gas mix cylinder into “Gas 1 “ port.
Make sure to use a demand flow regulator

## Slide 31
SDM-3R: GX-3R Placement
Plug in power adaptor and turn on the SDM-3R
Make sure the GX-3R is powered off before installing into the SDM-3R
When placing the GX-3R into the docking station, remember to have the yellow tab engaged upwards

## Slide 32
Turn on
If an automatic test is set up to occur, it will start as soon as the GX-3R is connected to the SDM-3R
Close the red door completely, the GX-3R will turn on automatically
When the GX-3R is connected to the SDM-3R, the GX-3R’s date and time are automatically updated to the SDM-3R’s current time and date

## Slide 33
Bump Test
Press the BUMP button to start bump test
Note: Verify that the appropriate calibration gas cylinder is connected to the gas 1 fitting on the left side of the SDM-3R

## Slide 34
Bump Test
During the bump test, the GX-3R’s screen will display the current gas readings

## Slide 35
Bump Test
Once a bump test is complete, the LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor.
The “BUMP” LED light on the SDM-3R will be green when a bump test passes.
The GX-3R will shut off automatically after successful bump test.

## Slide 36
Bump Test
If the bump test fails, ”BUMP” LED light is red. The SDM-3R will then automatically perform a calibration
The LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor for both bump test and calibration
If calibration is successful, the GX-3R will automatically power off

## Slide 37
Calibration
Press the CAL button until you hear the pump to start calibration
Make sure calibration gas concentrations match values shown on display

## Slide 38
Calibration
During calibration, the display shows current gas readings.
Calibration will take approximately 90 seconds.

## Slide 39
Calibration
Once calibration is complete, the LCD screen will alternate between the gas concentration and the letter “P” for pass or an “F” for fail for each sensor
The “CAL” LED light on the SDM-3R will be green when calibration is passed
“CAL” LED light is red if there if a failure. The GX-3R will power off after a successful calibration

## Slide 40
Removing from SDM-3R
Release the tab on the red door
Remove the GX-3R

## Slide 41
Data Transfer
Insert a USB flash drive with unit powered on.
To download the calibration and bump test data, press and hold COPY until the COPY LED is solid red, then release.
To download stored data from the gas monitor, keep pressing the COPY button until LED is flashing orange.
The USB flash drive can be removed when the COPY LED flashes green

## Slide 42
Unscrew the captive screws on the GX-3R.
Sensor/Filter replacement

## Slide 43
Sensor/Filter replacement
O2 Sensor:
ESR-X13P
LEL Sensor:
NCR-6309
H2S/CO:
ESR-A1DP
Sensor Dust Filter:
33-0181-10 (10 pk)
Scrubber Disk Filter for H2S/CO:
33-7130 (5 pk)
Scrubber Disk Filter for LEL:
33-7131 (5 pk)

## Slide 44
Questions?