## Slide 1
GX-2012 Confined Space / Multi-Gas Monitor
User Training

## Slide 2
Overview
Notes: Explain risks of potential cracked filter holder

## Slide 3
Inlet / Exhaust
Notes: Explain risks of potential cracked filter holder

## Slide 4
Filter Replacement
Cotton
Hydrophobic Filter
Inlet

## Slide 5
Damaged LCD Display
Replace the LCD display when it is damaged
Missing characters
Blotches

## Slide 6
LED Alarms
Alarm Lamp 1
Alarm Lamp 2
Alarm Lamp 3

## Slide 7
Buzzer
Buzzer (more than 95dB)
Notes: Explain blockage risks due to soils, etc.

## Slide 8
5 Key Touchpad
Notes: Brief explanation of buttons

## Slide 9
Instrument Serial Number
Side
Battery
Motherboard
Notes: Three locations for serial number

## Slide 10
Battery
Battery case

## Slide 11
Battery
Charging terminals
(Rechargeable)
Notes: Risk of oxidation and/or coated connections

## Slide 12
Battery
Battery release

## Slide 13
Battery
Alkaline
Rechargeable
Notes: Brief description of battery options

## Slide 14
Sensor Cover
Notes: Explain the risk of missing screws

## Slide 15
Hose and Probe
3.5” and 10’ ft hose available
Notes: Explanation of Hose and Probe. Inspect for cut or damaged hose, cracked probe housing.
How to attach to each other and to instrument

## Slide 16
GX-2012 Operation

## Slide 17
Press and hold the POWER ENTER button until you hear one beep, and release. The pump will engage.
Turning unit on

## Slide 18
Cal and Bump Reminder
Calibration due
Bump test due

## Slide 19
Start Up Display
Next Calibration
Next Bump Test
Date
Battery Level

## Slide 20
Start Up Display cont’d
Sensor(s) Installed
Full Scale
Warning Alarm

## Slide 21
Start Up Display cont’d
High Alarm
STEL
TWA

## Slide 22
LCD Overview
Peak level bar display
Monitors
- LEL / %volume CH4
- O2 - CO - H2S
Battery level
Blinking heart (microprocessor status)
Pump activity
Time indicator
Combustible display
automatically switches - %LEL / %volume

## Slide 23
Display Mode
In normal measuring mode press DISPLAY(ADJ)

## Slide 24
Display Mode
PEAK
STEL
TWA
Continue to press DISPLAY(ADJ) to scroll through the modes

## Slide 25
Alarm point
Operating time
Date
Continue to press DISPLAY(ADJ) to scroll through the modes
Display Mode

## Slide 26
Remaining log time
Clear datalogger
User ID
Continue to press DISPLAY(ADJ) to scroll through the modes
Display Mode

## Slide 27
Station ID
Snap logging
Peak bar select
Continue to press DISPLAY(ADJ) to scroll through the modes
Display Mode

## Slide 28
Changing Operating Modes
In normal measuring mode press and hold (SHIFT) for 4 seconds

## Slide 29
Normal Operating Mode
Leak Check Mode
Bar Hole Mode
Continue to press (SHIFT) to scroll through the modes
Changing Operating Modes

## Slide 30
Fresh Air Adjust
ALWAYS perform fresh air adjust in a clean, fresh air environment
In normal measuring mode press and hold AIR▲

## Slide 31
Keep holding AIR▲
Release AIR▲
Screen returns to normal measuring mode
Fresh Air Adjust

## Slide 32
Auto Range % LEL - % Vol.
As % LEL rises above 100%, reading automatically switches to % Vol.

## Slide 33
Pump Stall
The pump will automatically pause. A “FAIL” and a “LOW FLOW” status show on the screen

## Slide 34
Press and hold the POWER ENTER button until the pump disengages.
Turning unit off

## Slide 35
GX-2012 Bump Settings

## Slide 36
Press and hold the AIR▲ and (SHIFT)▼ buttons, then press and hold the POWER ENTER button.
When you hear a beep, release the buttons. The unit will prompt you for a password.
Maintenance Menu

## Slide 37
Password is 2102
Enter the password by using the AIR▲ and (SHIFT)▼ buttons to select each password number and then pressing and releasing POWER ENTER to enter it and move on to the next one.
Password

## Slide 38
Arrow down (SHIFT)▼ to LATCHING and press ENTER.
Latching

## Slide 39
Set to “OFF” using AIR▲ or (SHIFT)▼ then press ENTER
Latching

## Slide 40
Arrow down (SHIFT)▼ to BUMP DSP and press ENTER.
Bump Display

## Slide 41
Set to “ON” using AIR▲ or (SHIFT)▼ then press ENTER
Bump Display

## Slide 42
Arrow down (SHIFT)▼ to BUMP SET and press ENTER.
Bump Set

## Slide 43
Scroll down (SHIFT)▼ the list to “AUTO CAL” and press ENTER
Change to OFF
Bump Set

## Slide 44
Arrow down (SHIFT)▼ to ESCAPE and press ENTER.
Bump Set

## Slide 45
Arrow down (SHIFT)▼ to BP INTVL and press ENTER.
Bump Interval

## Slide 46
Set to 1 day not 30 by using (SHIFT)▼
Press ENTER
Bump Interval

## Slide 47
Arrow down (SHIFT)▼ to BP RMNDR and press ENTER.
Bump Reminder

## Slide 48
Set to “ON” using AIR▲ or (SHIFT)▼
Press ENTER
Bump Reminder

## Slide 49
Arrow down (SHIFT)▼ to AUTOZERO and press ENTER.
Auto Zero

## Slide 50
Set to “ON” using AIR▲ or (SHIFT)▼
Press ENTER
Auto Zero

## Slide 51
GX-2012 Bump Procedure

## Slide 52
When bump is due press and hold the (SHIFT)▼ button, then press the DISPLAY(ADJ) button and release both.
Bump

## Slide 53
AIR CAL (fresh air adjust) is the first menu to pop up and it is recommended to fresh air adjust the unit before performing a bump test.
Press POWER ENTER.
Bump

## Slide 54
Push and hold AIR▲ until the screen prompts you to release.
Bump

## Slide 55
Push AIR▲ or (SHIFT)▼ to scroll to BUMP and press POWER ENTER
Bump

## Slide 56
Push AIR▲ or (SHIFT)▼ to scroll to BUMP and press POWER ENTER
Bump

## Slide 57
Connect tubing from the demand flow regulator to the inlet fitting of the unit
Bump

## Slide 58
Press POWER ENTER.
The unit will countdown from 30 seconds, and the values will begin to flash.
Bump

## Slide 59
After 30 seconds, the unit will show the PASS/FAIL status.
If the bump test failed, calibration or service may be required
Remove tubing from inlet and press POWER ENTER to acknowledge the results.
Bump

## Slide 60
Bump
Push AIR▲ or (SHIFT)▼ to scroll to NORMAL and press POWER ENTER.
Unit will return to Normal measuring mode

## Slide 61
Press and hold the CAL button until the CAL LED turns on then release it.
  During calibration, display will show current gas readings
SDM-2012 Calibration
If a % volume sensor is installed the station will calibrate using Gas 2 port.

## Slide 62
Questions?