## Slide 1: GX-2009
Service Program
Rev. 2/3/09

## Slide 2: GX-2009 Features & Benefits
  Simultaneously monitors for:
    Combustible gas
    Oxygen
    Carbon monoxide
    Hydrogen sulfide
  Audible, Visual and Vibratory alarms.
    Distinctive alarm sound
  Built-in data logging
    8 Alarm trends
    Log time from 10 to 300 hours
    100 calibration records

## Slide 3: GX-2009 Features & Benefits
STEL/TWA and Peak hold function
Time and Date
RFI shielded
Rotatable alligator clip for hands-free monitoring
IP-67 rating for dust and water resistance Rechargeable NiMH batteries
  Up to 20 hours of operation in non alarm condition.
3 hours for a full charge
RF shielded high impact plastic case

## Slide 4: Startup Procedure
Press and briefly hold the POWER/MODE button
  Release when you hear a beep
If Cal. Limit is on and depending on how it is set, you may see the following:
  CAL C--LIMIT (confirm to use)
  FAIL C--LIMIT (can’t use)
If Cal. Limit Check is set to None, then the date and time and days to next calibration will be displayed.

## Slide 5: Startup Continued
Date and Time will be displayed.
Battery voltage is displayed with alarm logic.
  AL--H (Latching alarms)
  AL--A (Self resetting alarms)
Full scale values for all channels
Warning alarm set point for all channels
High alarm set point for all channels
STEL alarm set point for CO and H2S channels
TWA alarm set point for CO and H2S channels

## Slide 6: Startup Continued
If a sensor should fail during startup:
  FAIL <sensor type> SENSOR will be displayed.
  Replace sensor and calibrate as required.
If Auto Zero Adjustment is set to on, the GX-2009 will perform an automatic fresh air adjustment. Display will indicate:
  Auto ZERO
Once completed the GX-2009 will now display current readings of sensors installed.

## Slide 7: Fresh Air Adjustment
Take the GX-2009 to an environment free of toxic or combustible gases and of normal oxygen.
  Press and hold the AIR button. The display will indicate: Air CAL HOLD AIR
  Release the AIR button when ADJ RELEASE is displayed.
  The GX-2009 should indicate: 0% LEL combustibles, 0 ppm CO, 0.0 ppm H2S and 20.9% Oxygen.

## Slide 8: Display Mode
Pressing the POWER MODE button access Display Mode items
  PEAK (The GX-2009 will display the peak readings for combustibles, H2S and CO and the lowest reading for oxygen since the instrument was turned on)
    To clear the peak readings, press and hold the AIR button. The GX-2009 will indicate HOLD AIR, CLEAr RELEASE. Release AIR button. Press and release the POWER MODE button to move to the next screen.

## Slide 9: Display Mode
Display Mode Selections
  STEL (Short Term Exposure Limit, for H2S and CO)
    Press the POWER MODE button to move to the next screen
  TWA (Time Weighted Average, for H2S and CO)
    Press the POWER MODE button to move to the next screen
  Full Scale Screen (Detection range full scale value for each channel is displayed)
    Press the POWER MODE button to return to the measuring mode.

## Slide 10: Calibration Mode
The Calibration Mode provides the following menu:
  Set the date and time
  Perform a fresh air adjustment
  Perform a span adjustment (all channels simultaneously- AUTO CAL)
  Perform a span adjustment (one channel at a time - ONE CAL)
  Discharge the batteries with the REFRESH function.

## Slide 11: Accessing Calibration Mode
Begin with the GX-2009 off.
Press and hold the AIR button, then press and hold the POWER MODE button. When you hear a beep, release the buttons.
Using the AIR button, scroll to the AUTO CAL screen.
Press the POWER MODE button.

## Slide 12: Auto Calibration
Auto Calibration adjusts all active sensors simultaneously.
  Using the AIR button, scroll to AUTO CAL, then press the POWER button.
  All calibration values will be displayed:
    CH4 50% LEL
    OXY 12% Vol
    CO 50 ppm
    H2S 25.0 ppm

## Slide 13: Auto Calibration
If the calibration values displayed do not match the gas cylinder values:
  Press and hold the AIR button, then press the POWER button, CH4 % LEL --- will be displayed.
  Press the POWER button, 50% LEL will be displayed and blinking.
  Press the AIR button to increase the value or press the AIR button and hold then press the POWER button to decrease the value.

## Slide 14: Auto Calibration
If the calibration values do not match the gas cylinder values:
  Once set, press the POWER button.
  Press the AIR button to move to the next gas. Repeat as necessary.
  Press POWER button when ESCAPE is displayed and you are back to the AUTO CAL screen.

## Slide 15: Auto Calibration
Applying gas to the GX-2009
  Press the POWER button to begin Auto Calibration.
  Attach regulator to top of calibration gas cylinder.
  Attach tubing with calibration cap to regulator.
  Attach the calibration adapter plate onto the GX-2009’s sensor face and turn on the regulator.
  Allow gas readings to stabilize or a maximum of two minutes.
  Press and release the POWER button.
  The GX-2009 will attempt to calibrate all four sensors. If calibration is successful, PASS will be displayed. If not FAIL will be displayed.

## Slide 16: Auto Calibration
Auto Cal PASS
  Turn the gas regulator off.
  Remove the calibration adapter plate.
  The GX-2009 will indicate AUTO CAL.
  Press the AIR button until START is displayed then press the POWER button to return to normal operation.

## Slide 17: Maintenance

## Slide 18: Alligator Clip Removal
Unscrew the three Phillips screws to remove the Alligator Clip.
Replacement Alligator Clip: P/N 13-0116RK.

## Slide 19: Disassembly
Turn the GX-2009 off.
Using a small Phillips screwdriver remove the four screws securing the case halves together.
Grasp the case and separate.

## Slide 20: Disassembly
Internal Components
  Sensor PCB
  NiMh batteries
  Main PCB
  Buzzer

## Slide 21: Front Case & Components
Body case, front: P/N 21-1880RK
Gasket for buzzer hole: P/N 07-6018RK
Vibration motor: P/N 30-1054RK
Flexible PCB: P/N 24-5000RK
LCD Assembly: P/N 51-1119RK

## Slide 22: Rear Case
Body case, rear P/N 21-1881RK, buzzer included.
Gasket (10 pack) 07-6017RK

## Slide 23: Batteries
NiMh Batteries
  49-1609RK (2 ea)
  Average life: 2 years
  Simple to replace

## Slide 24: Sensor PCB Assy.
Sensor Case with PCB P/N 21-1883RK.
Sensor PCB connected to Main PCB by Flexible Printed Circuit Board P/N 24-5000RK- Take care to unplug flexible PCB before removing Sensor Case or damage to flexible PCB will result.

## Slide 25: Flexible PCB
Flexible PCB, P/N 24-5000RK connects Sensor case with PCB to Main circuit board.
There are sockets on both sides of the PCB that must be unplugged prior to removal.

## Slide 26: Main PCB Assembly
P/N 57-2052RK
Has 1.55 volt silver oxide battery to maintain memory
  Replacement battery P/N 49-1050RK

## Slide 27: LCD Assembly
LCD is keyed for easy installation
Ribbon cable has built in connector that is easily unplugged from the main PCB

## Slide 28: Sensors
LEL Sensor P/N NC-6264A
O2 Sensor P/N OS-BM2
CO Sensor P/N ES-1821
H2S Sensor P/N ES-1827
Average life: 2 Years

## Slide 29: Filter Maintenance
There are filters used to scrub the sample to protect the CO and LEL sensors
  HC sensor filter P/N 33-7114RK
  CO sensor filter P/N 33-7102RK
  Filter life dependent on exposure to H2S

## Slide 30: Filter Maintenance
Filters
  H2S filter for HC sensor P/N 33-7114RK
  Charcoal filter for CO sensor P/N 33-7102RK
  Waterproof sensor cover P/N 33-0172RK

## Slide 31: User Setup Mode
Set Date & Time
Perform fresh air adjust
Perform span adjustment
Set alarm points
Discharge batteries (Refresh)
Turn password feature on or off & define password

## Slide 32: Accessing User Setup Mode
Press and hold the AIR button, then press and hold the POWER MODE button. Continue to hold buttons.
After three seconds you will hear a second beep, release both buttons at this time.

## Slide 33: Using User Setup Mode
If password is set to on, then 0000 PASSWORD will be displayed
  If password is set, then enter password at this time
  If you do not know the password or need to enter this mode, then use the password key 9002
If password is off, then DATE will be displayed

## Slide 34: Using the User Setup Mode
If you need to change the Date & Time, press the POWER MODE button to enter this menu, if not press the AIR button
To perform an AIR CAL, press the POWER MODE button, otherwise press the AIR button
To perform a AUTO CAL, press the POWER MODE button to enter this menu, otherwise, press the AIR button
To perform a ONE CAL, press the POWER MODE button to enter this menu, otherwise press the AIR button

## Slide 35: Using the Setup Mode
To change the ALARM Points, press the POWER MODE button to enter this menu, otherwise press the AIR button
To fully discharge the batteries and perform a REFRESH, press the POWER MODE button, otherwise press the AIR button.
To change the PASSWORD, press the POWER MODE button, otherwise press the AIR button. Press the POWER MODE button when START is displayed to return to normal operation.

## Slide 36: Factory Mode
Press and hold the AIR button, then press and hold the POWER MODE button. Continue to hold buttons.
After three seconds you will hear a second beep, continue to hold the buttons until you hear the third beep.
Release the buttons at this time.

## Slide 37: Factory Mode
The display will indicate 0000
  Enter PASSWORD 1994
The display will indicate DATE, press the POWER MODE button to enter, otherwise press the AIR button.

## Slide 38: Factory Mode
GAS COMB will be displayed. Press the POWER MODE button to enter.
  Gas Combinations can be selected
  You can select CH4 % LEL or HC % LEL
  Using the AIR button you can select each detector and turn it on or off
  Press the POWER MODE button when ESCAPE is displayed

## Slide 39: Factory Mode
Press the AIR button and ALARM-P will be displayed. To enter press the POWER MODE button, otherwise press the AIR button.
LATCHING will be displayed. Press the POWER MODE button to enter this menu. Press the AIR button to select on or off.

## Slide 40: Factory Mode
AIR CAL is now displayed. Press the POWER MODE to select otherwise, press the AIR button to move to the next menu.
AUTO CAL is now displayed. Press the POWER MODE to select otherwise, press the AIR button to move to the next menu.

## Slide 41: Factory Mode
ONE CAL is now displayed. Press the POWER MODE to select otherwise, press the AIR button to move to the next menu.
ROM / SUM is now displayed.
  Press the POWER MODE button to display 03281 E61R depending on firmware version. Press the AIR button.

## Slide 42: Factory Mode
A/D VALUE will be displayed. Press the POWER MODE button to view. Current A/D values will be displayed at this time. Press the POWER MODE button to END. Press the AIR button to move to the next screen.
REFRESH will be displayed. Press the POWER MODE button to enter or press the AIR button to move to the next screen.

## Slide 43: Factory Mode
DEFAULT will be displayed
  Default resets this instrument to original configuration. Only do this when absolutely necessary
Press the AIR button and START will be displayed. Press the POWER MODE button to restart the GX-2009 and return to normal operation

## Slide 44: Questions?