## Slide 1
SDM-6000
Bump and Calibration
LEL / O2 / H2S / CO with Benzene & High Range PID

## Slide 2
Turn on the SDM-6000 using the ON/OFF toggle switch on the side
Setup

## Slide 3
To set up the gas inlets, go to CONFIG on the function menu
Setup

## Slide 4
Tap on the INLET icon
Setup

## Slide 5
Make sure the concentrations on the screen matches your cylinders, if not then make appropriate changes by clicking on the gas name and the arrow on the right side. This will bring up a menu where you can change the gas concentrations
Take note of the gas inlet number and make sure “Actual Gas” and “Active” are both checked
Setup

## Slide 6
Scroll down using the down button to GAS 2
Make sure that the cylinder you are using matches the concentration on GAS 2, and make sure “Active” and “Actual Gas” are checked
Setup

## Slide 7
Tap OK to confirm
Tap the back button to return to the main function menu
Setup

## Slide 8
Insert the 5-Gas mixed cylinder into GAS 1 port, as was configured earlier.
Setup

## Slide 9
Setup
Insert the Benzene Cylinder into GAS 2 port as configured in the inlet setup

## Slide 10
Your setup should look like this.
Please make sure to use demand flow regulator as shown on the image.
Setup

## Slide 11
With the GX-6000 turned off, insert the unit into the docking station
Setup
Press the big red button.

## Slide 12
The status box as shown in the bottom of the picture, should have a (which means that the unit is connected)
A red background indicator as in the picture means that the unit is due for a calibration.
Setup

## Slide 13
Indicator Background Colors
Background Color | Meaning
White | If no icon is showing, indicates that corresponding SDM-6000 is operating normally. If icon is showing, indicates that corresponding SDM-6000 is operating normally and connected instrument is not due for calibration, inlet assignments are correct, and gas/in gas out cover is engaged
Orange | Calibration is due in 10 days or less on connected instrument
Red | Calibration is due on connected instrument
Yellow | Inlet assignment error for connected instrument
Dark Grey | Gas in/gas out cover not engaged on connected instrument

## Slide 14
Indicator Icons

## Slide 15
Tap on the BUMP icon to perform a bump test.
Bump
Unit will begin with “AIR Flush”

## Slide 16
Tap on the UNIT tab to view actual gas readings.
Bump

## Slide 17
Unit will go through: -AIR Flush -GAS1 Bump -GAS2 Bump -AIR Purge -and lastly the Result
Bump

## Slide 18
When the BUMP passes, you will get a green square on the status bar on the bottom left.
The unit will turn off automatically after a set amount of time (depends on your preferred settings) or you can go to the function tab and manually end the bump test by pressing the bump end button
Bump

## Slide 19
Once the bump test is done, go into the UNIT tab to make sure that the detector is off by pressing the ON/OFF action button on the lower left hand side. (If the unit is not off already)
Detector Off: Means to turn off the detector (this means it’s on)
Detector On: This will turn the detector on (this means the unit is initially off)
*It’s an action button
Bump

## Slide 20
Once the unit is off, you can remove the GX-6000 out of the docking station and put another unit in.
Bump

## Slide 21
With the GX-6000 turned off, install it into the docking station and press the red button on top.
To perform a calibration, tap on the CAL icon
Calibration

## Slide 22
The unit will perform a calibration and should you need to cancel the calibration you can just tap on the CAL Cancel icon anytime
Calibration

## Slide 23
This is what the screen will look like if some of the sensors fail. You will see red highlight on the failed sensors.
Ex: the O2 and CO failed
This screen shows a successful calibration. All of the sensors that have passed are highlighted green
Calibration
FAILED
PASSED

## Slide 24
Tube Calibration
GX-6000
Benzene

## Slide 25
Once it passes the Auto Cal in the dock station, perform a tube calibration.
To get in to calibration mode press and hold SHIFT+DISPLAY while in normal mode
Tube Calibration - Benzene

## Slide 26
It will prompt you to install tube holder and press POWER/ENTER to confirm.
Tube Calibration - Benzene

## Slide 27
Choose Auto Cal on the menu and press POWER/ENTER.
Tube Calibration - Benzene

## Slide 28
Choose concentration and press POWER/ENTER.
Scroll down until you see BNZ (Benzene) 5 ppm.
Tube Calibration - Benzene
(Make sure it matches your cylinder) and press POWER/ENTER to calibrate that channel.
Get your cylinder and demand flow regulator ready right after you press the POWER/ENTER button

## Slide 29
Once you press POWER/ENTER the monitor will countdown.
Apply the gas in to the meter.
Tube Calibration - Benzene
Wait for the meter to countdown and either you’ll get a pass or a fail.