Quick Reference Guide

#### Beacon 410 Series Programming

TOOLS REQUIRED:

None.

##### Configuration Menu

- Press and HOLD the ESCAPE and ENTER buttons.

  - Select Channel 1-4.

    - Press the UP or DN button to select the channel that you want to configure, then press the ENTER button.

  - Channel (x) Active.

    - Press the UP or DN button to select Not Used or Stand By then Press the ENTER button.

  - Channel (x) Alarm-1 Level.

    - Press the UP or DN button to change, then press the ENTER button.

  - Channel (x) Alarm-1 INC/DEC.

    - Press the UP or DN button to select INC or DEC then press the ENTER button.

  - Channel (x) Alarm-1 Latching.

    - Press the UP or DN button to select Latching or Self Resetting then press ENTER.

  - Channel (x) Alarm-1 On Delay. 1 Seconds

    - Press the UP or DN button to select from 0 seconds to 60 Seconds then press the ENTER button.

  - Channel (x) Alarm-2 Level.

    - Press UP or DN button to select then press the ENTER button

  - Channel (x) Alarm-2 INC/DEC

    - Press the UP or DN button to select INC or DEC then press the ENTER button.

  - Channel (x) Alarm-2 LATCHING.

    - Press the UP or DN button to select Latching or Self Resetting then press the ENTER button.

  - Channel (x) Alarm-2 On Delay.

    - Press the UP or DN button to select from 0 to 60 seconds then press the ENTER button.

  - Channel (x) Alarm-3 Level.

    - Press the UP or DN button to set then press ENTER button.

  - Channel (x) Alarm-3 INC/DEC

    - Press the UP or DN button to select INC or DEC then press the ENTER button.

  - Channel (x) Alarm Latching.

    - Press the UP or DN button to select Latching or Self Resetting then press the ENTER button.

  - Channel (x) Alarm-3 On Delay.

    - Press the UP or DN button to set On Delay from 0 to 60 seconds.  Press the ENTER button.

  - Channel (x) Relay -1 Assignment, Alarm-1

    - Press UP or DN button to select Alarm-1, Alarm-2, Alarm-3, Alarms 1 or 2, Alarms 1 or 3, Alarms 2 or 3, Alarms 1, 2, or 3 or Use As Channel Fail, then press ENTER button.

  - Channel (x) Relay-2 Assignment, Alarm-2

    - Press the UP or DN button to select Alarm-2, Alarm-3, Alarms 1 or 2, Alarms 1 or 3, Alarms 2 or 3, Alarms 1, 2 or 3, or Use As Channel Fail, then press ENTER button.

  - Channel (x) Relay-A Assignment, Alarm-3

    - Press the UP or DN button to select Alarms 1 or 2, Alarms 1 or 3, Alarms 2 or 3, Alarms 1, 2 or 3 or Use As Channel Fail for Alarm-1, Alarm-2 or Alarm-3 then press ENTER button.

  - Channel (x) Noise Filter, 2 Seconds.

    - Press the UP or DN button to select from 1 Second to 10 second then press Enter button.

  - Channel (x) Deadband

    - Press the UP or DN button to select from 0 to ??? then press the ENTER button.

  - Save Settings?

    - Y/N

    - Press the Y button

    - Settings Saved

    - Do you want more channels?

    - Press Y or N

Global Menu

- Press and hold the ESCAPE and UP/YES buttons to enter into the Global Menu

  - Continue Y/N, press the YES button

  - Strobe Installation

    - No Strobe Installed (displayed)

    - Press the UP or DN button to select No Strobe Installed or Strobe is Installed then press ENTER button.

  - Normal Relay States

    - Norm De-Energized (displayed)

    - Press the UP or DN button to select Norm Energized or Norm De-Energized, then press the ENTER button.

  - Relays ABCD Used As:

    - Channel Alarm Relays (displayed)

    - Press the UP or DN button to select Channel Alarm Relays or Common Alarm Relays, then press the ENTER button.

  - Buzzer Silence

    - Can Silence Buzzer (displayed)

    - Press UP or DN button to select Can’t Silence Buzer or Can Silence Buzzer, then press the ENTER button.

  - Save Settings?

    - Y/N

    - Press the YES button to save the settings

Mod Bus Menu

- Press the DN/NO and ENTER buttons simultaneously to enter ModBus Menu

- Continue Y/N?, press the YES button to continue.

  - ModBus Feature is Disabled (displayed)

    - Press the UP or DN button to select Inabled or Disabled, then press the ENTER button.

  - Slave ID: 1 (displayed)

    - Press the UP or DN button to select between 1 and 247 then press the ENTER button.

  - Baud Rate: 9600 (displayed)

    - Press the UP or DN button to select 14,400, 19,200, 4,800, 2,400 or 1,200, then press the ENTER button.

  - Parity: Even

    - Press the UP or DN button to select NONE, ODD or EVEN, then press the ENTER button.

  - Response Delay: 0 mS (displayed)

    - Press the UP or DN button to select from 0 to 20 mS then press the ENTER button.

  - Save Settings? Y/N

    - Press the YES button to save settings.

    - Settings Saved

Calibration Mode

- Press the hold the UP/YES and ENTER buttons to enter Calibration Mode

  - Continue Y/N?

    - Press YES button to continue

  - Select Operation

    - Set Cal Time Out (displayed)

    - Press the UP or DN button to select Air Adjust, Perform Gas Adjust, View Max Spans, Exit Cal Menu.

    - Press the ENTER buttin to enter the menue you require.

  - Set Cal Time Out

    - Press ENTER button

    - Press the UP or DN button to select from 5 minutes to 240 minutes.

    - Press the ENTER button when completed.

    - Settings Saved will be displayed.

  - Perform Air Adjust

    - Press ENTER button

    - Select the channel you want to Air Adjust

    - Press the UP button to select 2, 3, 4, 1 and 2, 1 and 3, 1 and 4, 2 and 3, 2 and 4, 3 and 4, 1, 2 and 3, 1, 2 and 4, 1, 3 and 4, 2, 3 and 4, 1, 2, 3 and 4.

    - Note:  Air adjust will be used with Direct Connect sensors only.

    - Press the ENTER button

  - Perform Gas Adjust

    - Press ENTER button

    - Press the UP or DN button to select the same channels as above.

    - Note: Gas Adjust to be used on Direct Connect sensors only.

    - Press the ENTER button

  - View Max Spans