Gas Detection For Life

Eagle 1

Hands-on Worksheets

Objective:

To increase students ability to properly repair a standard Eagle portable gas detection instrument.

Tools Required:

Philips and Standard screwdrivers, digital volt meter and calibration kit.

1)	Customer claims that your instrument will not respond properly to 	methane 	gas.  What are your findings?

A)	Sensor Date Code:  ____________  List the age of your sensor 				based upon the date code found:  _______________________

B)	Is the sensor under warranty?   Yes / No  What is the standard 				warranty of the LEL sensor?  ___________ years.

C)	What is the span as found?  ____________

D)	What is the maximum span?  ____________

E)	Can the instrument be calibrated?  Yes / No

F)	What is the maximum time calibration gas should be applied to 				the Eagle?  __________ minutes.

G)	What position is the calibration switch on the main board set?

CH4 / HEX.

H)	What happens to the LEL sensor voltage if the switch is placed 				in the HEX position?  _______________

2)	Customer claims that the Oxygen sensor on their instrument will not set to 	12% 	O2.  What could be the cause?  (circle correct answers)

A)	Oxygen sensor output voltage too high.

B)	Incorrect gas being used.

C)	Leak in probe or inlet fitting.

D)	Inlet filter dirty.

3)	What is the output voltage of the O2 sensor in your monitor?  _________ 	mV

4)	What is the date code?  _________

5)	Using RKI Guideline GL-0004, when should the O2 sensor for the 	Eagle be 	replaced?   After ______ years.

6)	What is the flow rate on your Eagle?  ________ SCFH

7)	Where is the low flow alarm setting on your instrument?  _______ SCFH.

8)	What should the flow fault alarm be set to?  ________ SCFH +/- ______

9)	The charcoal filter in the Eagle flow block performs what function?  		_____________________________________________________

10)	Is it okay to leave the batteries out of an Eagle with a super toxic 	electrochemical sensor installed?  Yes  / No  Why?  	__________________________________________________

11)	You receive an Eagle that you are unable to enter into the calibration mode 	by pressing the SHIFT and DISPLAY buttons.  What could be wrong?   	______________________________________________

12)	Is it okay to use rubber tubing on your demand flow regulator when 	calibrating 	with a multi gas blend cylinder?  Yes / No

13)	You are asked to calibrate an Eagle and it fails when performing a 	AUTO CAL.  	What could be happening?   Circle correct answer.

A)	Cal gas settings incorrect

B)	One or more sensors are bad

C)	Leak in sample system

D)	Cylinder empty

E)	All of the above

F)	None of the above

14)	You have an Eagle that has a very unstable zero on the LEL sensor.  What 	could cause this problem?  Circle correct answer.

A)	Contaminated flow system (fuel or other flammable substance)

B)	Improper sensor voltage

C)	Leak in sample system

D)	Restrictive exhaust port

15)	The customer wants to change the calibration from methane to hexane.  Go into 	the setup mode and change the gas type and range.

16)	Calibrate your instrument to hexane.  Provide the as found / 	maximum and as 	left data below:

as found:  _____     maximum span:  _____   as left:  _____

17)	After calibrating to hexane, exit out of the calibration mode and apply 50% 	LEL 	methane to the instrument and record the response.   Reading:  _____

18)	Go back into the calibration mode and reset the gas type for methane 	and 	recalibrate.

19)	What items would you check if your Eagle went into flow fail after warm up?

A)	Probe filter

B)	Internal filter

C)	Pump

D)	DP Switch Setting

E)	Analog PCB

20)	If you have an Eagle that will not power up, what should you look for?

Circle all correct answers

A)	Missing or expired sensor

B)	Corrosion on battery contact terminals

C)	Damaged I/S board assembly

D)	Damaged or unplugged main ribbon cable

E)	Dead or missing batteries