Gas Detection For Life

Eagle 2

Hands-on Worksheets

Objective:

To increase students ability to properly repair a standard Eagle 2 portable gas detection instrument.

Tools Required:

Phillips and Standard screwdrivers, digital volt meter and calibration kit.

1)	How do you enter into the Set-up mode?

A)	Press the DISPLAY and AIR buttons and turn on

B)	Press the ENTER and RANGE buttons

C)	Press the AIR and RANGE buttons and turn on

D)	Press the DISPLAY and RANGE buttons and turn on

2)	Enter into the Set up mode and do the following:

A)	Set the Date and Time to current date and time

B)	Set the Date format to DD/MM/YYYY

C)	Set Relative Response Factor to ON

D) 	Exit out of Set up mode

3)	How do your enter into Factory Mode?

Press and hold what buttons then power the instrument up.

A)  	UP

B)	DOWN

C)	DISPLAY

4)	What is the password to enter into the Factory Mode?

1994

5)	You have an Eagle 2 that has low or no methane gas response but seems to work 	okay on hexane.  What should you look at?  	First check to verify if METHANE ELIMINATION MODE is on, if not then sensor needs to be replaced.

6)	An Eagle 2 is returned that will not charge, what do you suspect and what should 	be done to correct the problem?

CHECK THE BATTERIES / CHECK THE BATTERY PACK / FUSE OK?

7)	An Eagle 2 is received with low capacity Ni-MH batteries installed, but they 	seem to work okay.  Should you leave them in?  Yes  / No

Why?  BATTERIES COULD be damaged if attempting to charge, batteries could leak and cause damage to Eagle 2.  The charger is designed for use only with rKI supplied NiMH batteries.

8)	You have received an Eagle 2 from a customer that has a bad LEL sensor.  To 	repair the instrument you need to replace the sensor but you only have a NC-	6260A Eagle 1 sensor in stock, can you use this sensor in the Eagle 2?  Yes / No

9)	Using a flow meter, measure the flow on your Eagle 2.  What is the current flow 	rate?  1.6-2   SCFH.

10)	Restrict the flow using the valve on the flow meter and test the low flow alarm.  	Where is it set?  0.6 +/- 0.1  SCFH.

11)	When testing the Eagle 2 with 25 ppm H2S, you see the CO sensor start to 	respond.  What is the problem and what should be done to correct it?

Problem: CROSS INTERFERENCE  Solution:  CHANGE CHARCOAL FILTER

12)	What is the typical output voltage of the oxygen sensor used in the Eagle 2?

10-20 mV.

13)	Why do you need a CO2 scrubber when setting the fresh air zero on an Eagle 2 	with 0-10000 ppm CO2 range?

BECAUSE THERE IS ABOUT 300-400PPM CO2 IN STANDARD, FRESH AIR IN AN OFFICE ENVIRONMENT

14)	You are calibrating an Eagle 2 and find that the oxygen fails to set at 12% by 	volume.  What is the most likely cause of the problem and what should you 	check?

LEAK IN FLOW SYSTEM (INCLUDING PROBE)

15)	You are calibrating an Eagle 2 with high range PID, what is the calibration gas 	type and value you should be using?

ISOBUTYLENE  gas    100  PPM