| Features: | Fast warm-up time |  |  |  |
| --- | --- | --- | --- | --- |
|  | Good zero stability |  | #### Part Number: | ES-23SD-H2SE |
|  | Quick response time |  | #### Sensor Application: | GD-K8A |

| #### Technical Specifications |  |  |  |  |
| --- | --- | --- | --- | --- |
| ##### Measuring Principle | Amperometric 3-electrode sensor |  | #### Accuracy | +/- 10 % of reading |
| Range of Measurement | 0 – 0.20 ppm |  | Repeatability | +/- 5% of reading |
| Resolution | 1% of full scale |  | T90 Response time (20°C, 2 min. exposure) | 90 seconds |

| Operating Conditions |  |  |  |  |
| --- | --- | --- | --- | --- |
| Temperature Range | -20°C to +45°C |  | Life Expectancy | 2-3 Years |
| Humidity Range | 10-95% RH, Non Condensing |  | Warranty | 1 Year |
| ###### Known Gas Interferences |  |  |  |  |

| Gas | PPM Gas Applied | Reading |  | Gas | PPM Gas Applied | Reading |
| --- | --- | --- | --- | --- | --- | --- |
| Acetone ((CH3)2CO) | 1% vol | 0.05 ppm |  | Hydrogen Chloride (HCl) | 0.26 ppm | 0.05 ppm |
| Ammonia (NH3) | 310 ppm | 0.05 ppm |  | Hydrogen Sulfide (H2S) | 0.056 ppm | 0.05 ppm |
| Arsine (AsH3) | 0.11 ppm | 0.05 ppm |  | Isopropyl Alcohol (IPA) ((CH3)2CHOH) | 1% vol | 0.05 ppm |
| Carbon Monoxide (CO) | 0.5% vol | 0.05 ppm |  | Nitric Oxide (NO) | 160 ppm | -0.05 ppm |
| Chlorine (Cl2) | 0.16 ppm | -0.05 ppm |  | Nitrogen Dioxide (NO2) | 0.21 ppm | -0.05 ppm |
| Chlorine Trifluoride (ClF3) |  |  |  | Ozone (O3) | 0.23 ppm | -0.05 ppm |
| Diborane (B2H6) | 3.2 ppm | 0.05 ppm |  | Phosphine (Ph3) | 0.08 ppm | 0.05 ppm |
| Disilane (Si2H6) | 0.72 ppm | 0.05 ppm |  | Silane (SiH4) | 1.5 ppm | 0.05 ppm |
| Germanium Tetrahydride (GeH4) | 0.62 ppm | 0.05 ppm |  | Sulfur Dioxide (SO2) | 1.8 ppm | 0.05 ppm |
| Hydrogen (H2) | 0.74% vol | 0.05 ppm |  |  |  |  |