Gas Detection For Life

GX-2009

Hands-on Worksheets

Objective:

To increase students ability to properly troubleshoot and repair the GX-2009.

Tools Required:

Small Philips and Standard screwdrivers, digital volt meter and calibration kit.

1)	A GX-2009 is returned for service and the complaint is that when charging the 	instrument the red charge light does not come on when the instrument is put into 	the charger.  What would you do to troubleshoot this problem?

________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

2)	What will happen to the GX-2009 if the user performs a "Default?"

____________________________________________________________________________________________________________________________________________

3)	You need to calibrate a GX-2009 and you only have test gas with the following 	components, 50% LEL Methane, 50 ppm CO, 25 ppm H2S, 20.9% O2 balance 	N2.  Can you perform a Auto Cal?  ___________________________________________________________________

4)	You have a GX-2009 that the customer complains that the instrument will fail 	calibration on combustibles channel.  How would you go about troubleshooting 	this?

________________________________________________________________	_____________________________________________________________________

5)	You get an OS-BM2 oxygen sensor back from a customer and measure the 	output and it is 300mV.  Is this okay?  Yes  /  No   Why?  	________________________

6)	You receive a 2009 back for repair and the instrument will not power up even with freshly charged batteries.  What would you need to replace? __________________________________________________________________

7)	Customer returns a GX-2009 with the comment that it fails H2S upon start-up.  You open up the unit and see that the sensor gold contacts are silver and green.  What has happened?  ________________________________

8)	You replace the H2S sensor and it still fails, what is the most likely cause of the problem and what has to be done to correct the fault? ____________________________________________________________________________________________________________________________________________

9)	You receive a GX-2009 and the customer wants it calibrate to Hexane on the LEL channel.  What do you need to do before calibrating?  ______________________________________________________________________

10)	Why does is the CO sensor scrubber black and what is its purpose?  ______________________________________________________________________