## Slide 1: Verifying and Replacing Fuse in an Eagle 2 Battery Housing
February 10, 2016

## Slide 2: Parts Required
Soldering iron
Solder
Solder braid
Small Vice
Fuse P/N 43-4142RK-10
Digital Multimeter with quick hook clip leads

## Slide 3: Remove pack and batteries from Eagle 2

## Slide 4: Setup to Test Fuse
Need DMM set to measure diode function

## Slide 5: Measuring Points
Measuring three diodes through the 1.5 A fuse in the battery housing. CN1 on left and EC 1 or EC 2 on right.

## Slide 6: Open Fuse
Place DMM into Diode Checking Mode. Attach positive DMM lead to CN1 lead,
and attach or touch black DMM lead to EC1 or EC2. If OL is displayed, fuse is open.

## Slide 7: Good Fuse
If the fuse is good, you will see the diode drop across the three Schottky
Diodes, Fuse F1 and BC1 which is connected to EC1 and EC2.

## Slide 8: Reverse leads to check diodes in box with good fuse

## Slide 9: Procedure for replacing Fuse

## Slide 10: Using a Phillips screwdriver, remove the two screws that connect the PCB to housing.

## Slide 11: Plug in charger cable to use as a heat sink. Do not plug the charger into power outlet!

## Slide 12: Carefully remove solder from charger socket. Solder braid works well here.

## Slide 13: Remove PCB from housing

## Slide 14: Remove thermistor from housing. Note, make sure thermistor is reinstalled after fuse replacement

## Slide 15: Using soldering iron, heat side of fuse and lift from PCB. Do this to both sides and remove the fuse.

## Slide 16: Once fuse is removed. Clean old solder from pads and add small amount of fresh solder

## Slide 17: Lay new fuse on pads and solder in place.

## Slide 18: New fuse installed!

## Slide 19: Place PCB back into housing

## Slide 20: Reattach charger socket to PCB. Keep charger plug in connector. Remove charger plug when completed.

## Slide 21: Replace Phillips screws.

## Slide 22: Replace batteries and verify that battery pack charges.

## Slide 23: Charge Module
Plug charger into AC outlet
  Green LED will illuminate, then Yellow LED then both Green and Yellow LEDs will remain lit.
  This is what you see when the charger boots up.
Plug charger cord into Eagle 2 battery pack
  The Green LED will go out and the Yellow will remain lit.
  If not, then there is no charging current flowing into batteries.
  Further investigation is required.