S-Series LEL Detector Calibration

Tools required: Digital Voltmeter and normal test leads, small straight slot screw driver, calibration kit.

- Put the RKI controller or the customer supplied PLC or DCS system into bypass.

- Unscrew the explosion proof cover on the S-Series head.

- Put the Digital Voltmeter on Millivolt range and plug the red lead into the red jack on the amplifier(round black board) then plug the black lead into the black jack.

- The Digital voltmeter should read 100 millivolts at this time if it does not read 100millivolts then the zero potentiometer (bottom right hand side) should be adjusted until the Digital Voltmeter reads 100 millivolts plus or minus 2 millivolts.

- Calibration cup should then be put on the sensor and the gas turned on.

- The formula to get the voltage that should be seen with the gas applied to the sensor is.

Gas value

Full Scale  X 400  + 100= Set point

So if this was an LEL sensor and the calibration gas value is 50% LEL then

50%LEL

100%LEL X 400 + 100=300millivolts

- With the gas applied to the sensor for 2 minutes the gas value should be 300 millivolts plus or minus 2 millivots if the above values are the same. If this is not correct the Span potentiometer (bottom left of amplifier) would then be adjusted until the proper reading is obtained on the Digital voltmeter.

- The calibration gas is then removed from the sensor and the cover placed back on the explosion proof housing.

Note: Do not go back and rezero the instrument right after the gas has been applied it will take a while for the instrument to go back to zero.