## Slide 1: Beacon 410
4 Channel Wall Mount Gas Detection Controller

## Slide 2: Beacon 410 Features
PCB Control Switches
ESCAPE
UP/YES
DOWN/NO
ENTER
Alarm LEDs
Alarm 1, yellow
Alarm 2, orange
Alarm 3, red
Fail, yellow
On/Off toggle switch
94 dB Buzzer Alarm
Alarm Reset Switch
Backlit LCD Display
4 Channels Displayed Simultaneously
NEMA-4X weather resistant enclosure
Accepts any 4-20 mA transmitter or direct connect sensors for LEL / O2 / H2S / CO / CO2
3 programmable alarm levels per channel

## Slide 3: Alarm Indications
Condition | Cause | Visual Indications | Audible Indications
Alarm 1 | Increasing (decreasing for O2) gas reading at or above the alarm 1 set point | Alarm 1 LED on Gas reading alternates with Alarm-1 message Strobe (if installed) activates | Pulsing Tone
Alarm 2 | Increasing gas reading at or above the alarm 2 set point | Alarm 2 LED is on Gas readings alternates with ALARM 1 and ALARM-2 messages Strobe (if installed) activates | Pulsing Tone
Fail | Disconnected or misconnected detector head wiring Display reading at -10% of full scale or lower Defective component | Fail LED is on Fail message replaces gas reading Strobe (if installed) activates | Pulsing Tone
Low Power | No AC power and DC power source (primary or backup) less than 21.5 volts. | Fail LED is on SUPPLY VOLTAGE IS TOO LOW LOW POWER STANDBY message and actual voltage of incoming DC power | None

## Slide 4: Beacon 410 Calibration
The Beacon 410 can support both direct connect (internal amp) and 4-20 mA transmitter (remote amp) detector heads.
• Direct connect detector heads only.
  If all the active channels are direct connect detector heads, then all calibration adjustments are made at the Beacon 410 after calibration gas is applied to the detectors.
• 4-20 mA transmitter detector heads only.
  If all the active channels are 4 - 20 mA transmitter detector heads, then all calibration adjustments are made at the detector head when calibration gas is applied to the detectors.

## Slide 5: Configuration Menu
Press the ESCAPE and ENTER buttons to enter into Configuration menu.
Press YES to continue.

## Slide 6: Configuration Menu
Select the channel you would like to view.
You can make the channel ACTIVE.

## Slide 7: Configuration Menu
If ACTIVE is selected, you can step through all parameters for this channel using the ENTER button.
Adjust parameters using the UP or DOWN button

## Slide 8: Channel Parameters
Parameters include:
  Alarm 1
    Alarm 1 LEVEL
    Alarm 1 INC/DEC
    Alarm 1 LATCHING
    Alarm 1 ON DELAY
  ALARM 2
    Alarm 2 LEVEL
    Alarm 2 INC/DEC
    Alarm 2 LATCHING
    Alarm 2 ON DELAY
  Alarm 3
    Alarm 3 LEVEL
    Alarm 3 INC/DEC
    Alarm 3 LATCHING
    Alarm 3 ON DELAY
  Relay 1 Assignment Alarm 1
  Relay 2 Assignment Alarm 2
  Noise Filter
  Deadband

## Slide 9: Configuration Mode
Channel can be placed in STANDBY mode.
All alarms are off line for this channel.
Sensor may be attached but not active.

## Slide 10: Configuration Mode
The channel can be set to NOT USED.
No sensor or transmitter would be wired to the panel.

## Slide 11: Calibration Mode
Press and hold the UP/YES and ENTER buttons to enter into the calibration mode.
Press YES to continue.

## Slide 12: Calibration Mode
Set Cal Timeout from 5 minutes to 240 minutes.
Press ENTER button to continue.

## Slide 13: Calibration Mode
Perform Air Adjust
Select channel you would like to Air adjust.
You can select any combination of the four sensors attached.

## Slide 14: Calibration Mode
When Perform Gas Adjust is displayed, press the ENTER button to continue.
Apply gas to the channels selected.

## Slide 15: Calibration Mode
For 4-20mA transmitters, calibration must be done at the amplifier.

## Slide 16: Calibration Mode
Gas reading will be displayed.

## Slide 17: Calibration Mode
The display will prompt you to apply calibration gas.
Once done with the calibration at the amplifier, press the ENTER button to continue.

## Slide 18: Calibration Mode
To view max spans, press and hold the ENTER button.

## Slide 19: Calibration Mode
Once completed, press the ENTER button to EXIT calibration mode.

## Slide 20: Entering Global Menu
Press and hold the ESCAPE and UP/YES buttons to enter into the global menu.

## Slide 21: Global Menu Settings
Strobe
  No Strobe Installed
  Strobe Installed
Normal Relay States:
  Norm DE-ENERGIZED
  Norm ENERGIZED
Relays ABCD Used As:
  Common Alarm Relays
  Channel Alarm Relays
Buzzer Silence
  Can Silence Buzzer
  Can’t Silence Buzzer

## Slide 22: Global Menu
After all Global parameters are set, save settings by pressing the Yes button.

## Slide 23: Entering MODBUS Mode
Press and hold the DOWN/NO and ENTER buttons to enter into the MODBUS menu.
Press YES to proceed.

## Slide 24: Modbus Settings
Modbus Feature
  Enabled
  Disabled
Modbus Setup
  Slave ID
  Use UP/Down to set
Baud Rate
  9600
  Use UP/Down to set
Parity
  Even
  None
  Odd
Response Delay
  0mS
  Use UP/Down to set
Save Settings
  Select Yes or No

## Slide 25: Input Set up Mode
With the Beacon 410 off, press and hold the ENTER button and turn the Beacon on.
Press the YES button to proceed.

## Slide 26: Input Setup
Select Channel (1-4)
  Preamp Direct
  Oxygen Direct (25%)
  RWGB Direct (LEL)
  4-20mA XMTR
Gas Setup is displayed.
Set for the proper gas type.
Save Settings
  Select Yes/No

## Slide 27: Questions?