## Slide 1: GD-70D series Solar Cell, Semiconductor & LCD Factories
Improved Function
Global Standard
Environmentally
Friendly Design

## Slide 2: Concept
Universal design
1) Basic components consists of:
* Main PCB, Sensor assembly and Pump unit
2) Easy maintenance from front access. No tools required.
3) Large LCD (White-color backlight)
4) No need for daily check. Auto flow adjustment and self diagnosis.
Compact and Renewable
5) Power-saving technology
6) Simple Operation
7) Smaller size and weight
8) Recyclable sensor/amplifier module
International Certificate
9) RoHS/CE marking
Global standard
Environmentally Friendly
Design
Improved Function
GD-70D has been developed to fulfill all the customer’s wants and needs

## Slide 3: Features for GD-70D
1. Common platform for all detection methods
2. Universal Main unit for all sensor types
3. Multifunctional intelligent Sensor unit
4. No internal tubing (Main unit) / No Coil (Pump unit)
5. Front access, no tools required for easy replacement of sensor & pump
6. Large size LCD (Easily viewable)
8. Reduced maintenance through enhanced firmware functions
9. Small Footprint
10. Simple upgrade from existing units
11. Environmentally Friendly
12. Global standard certificate
7. Various communication method available (4-20mA, NT and POE)

## Slide 4: Upgraded Function
GD-70D consists of three independent modules: main unit, sensor unit and pump unit.
a) Main unit
c) Pump Unit
b) Sensor unit
Pyrolyzer (PLU-70D)
Connecting pyrolyzer (PLU-70D)
to GD-70D, TEOS & NF3 can be
measured.
Temperature for PLU-70D is
automatically controlled.
Signal converter (SD-70SC)
Setting Signal converter the
(SD-70SC) & GD-70D can be
Networked.
1. Common platform for all detection methods

## Slide 5: Upgraded Function
Main unit can be used with all sensor units with various detection principles
Converting from SiH4 to H2 is simple!
Just remove the SiH4 sensor unit and replace with an H2 sensor unit.
Various network systems available:
GD-70D : 4-20 mA DC GD-70D-NT ：NT GD-70D-ET ：PoE
SiH4 Detector
H2 Detector
H2 Sensor unit (SGU)
SiH４ Sensor unit (ESU)
2. Universal Main unit
①Remove
②Set
The GD-70D has functions that will not allow the insertion of incorrect sensor units. Every sensor unit has a unique identification which prevents installation error.

## Slide 6: Upgraded Function
Electro Chemical
Semiconductor
Galvanic Cell
New Ceramic
Pyrolysis-Particle
There are five different detection technologies available.
Each sensor unit can monitor and control measuring status.
In the event of an alarm, its measuring status can be memorized by the sensor.
3. Multifunctional Intelligent Sensor Unit
Available Sensor Units
Each sensor has onboard CPU and provides multiple functions

## Slide 7: Upgraded Function
The diaphragm is built in Pump unit for GD-70D
The magnetic coil is built in to the Main Unit.
There are no tubes inside Main Unit to simplify maintenance
Magnet coil is built in Main unit.
Diaphragm is easy to replace and built in to the pump unit.
Main Unit
Pump Unit
4. No internal tubing and no pump coil to replace
Connection for
Pump unit
Connection for
Sensor unit

## Slide 8: Upgraded Function
Sensor Unit/Pump unit can be removed easily with one touch of button.
Main unit
Sensor unit
Pump unit
Wall mounting bracket
Remove and attach items with one touch of a button
5. Front access for easy maintenance

## Slide 9: Upgraded Function
GD-70D has a large easy-to-read LCD that displays all information such as measuring gas, gas concentration, flow level and connecting devices.
Dual display for gas concentration : Analog bar graph and digital readouts
Power lamp
Alarm lamps
Fault
Display for gas Concentration
Alarm point
Gas Concentration
Pyrolyzer connection
Communication
status
Maintenance status
Inhibit
(Offline: Communication for PC.)
Measuring gas
Units of Measure
Flow rate
Control switch
Name plate for Sensor unit
(automatic adjustment)
6. Large size LCD
Notes: 9

## Slide 10: Upgraded Function
・Automatic flow rate control
GD-70D can automatically control flow rate to preset flow level, which makes gas measuring stabilized.
Preset flow rate
Lower flow rate
Over flow rate
Flow UP
Flow DOWN
・Self-diagnosis function
GD-70D continuously monitors various measuring status in detail.
・Data logging functions
Event history, Alarm history, Calibration history, Alarm trend with 1 sec interval
Automatically adjusted
7. Reduced maintenance through enhanced troubleshooting firmware functions
Notes: 10

## Slide 11: Global Standard
RoHS
CE marking
IEC Standard
11. Global standard
Notes: 11

## Slide 12: Specifications
Main Unit |  |  | 
Model | GD-70D | GD-70D-NT | GD-70D-ET
Communication | 4-20mADC | DC power line communication | POE Method
Detection principle | Different type depending upon sensor unit and detectable gas (See table on back page for list of sensor types and detectable gases) |  | 
Sampling method | Sample-drawing（Auto-Adjustment of flow rate) |  | 
Display | Large LCD （White back light） |  | 
 | ・Gas Concentration |  | 
 | ・Flow rate / Communication status / Pyrolizer status / Character display |  | 
 | ・Measuring gas / Error code / Content of error / Character message display |  | 
Display | 1ST alarm： red / 2nd alarm： red / Fault alarm：yellow |  | 
External Output | 1st alarm / 2nd alarm / Trouble alarm ：Relay contact output for each alarm |  | 
Self diagnosis | ・System failure |  | 
 | ・Sensor failure |  | 
 | ・Flow failure |  | 
 | ・Communication failure(NT / ET) |  | 
Date logging function | ・Event history / Alarm history |  | 
 | ・Calibration history |  | 
 | ・Alarm trend （180 sec Before / after 1st alarm） |  | 
Operating conditions | 0-40℃ （Without rapid change） ・30-70%RH （non-condensing） |  | 
Operational Setting | All operational setting are user adjustable by using the front panel |  | 
Power requirement | DC 24V±10% Approx. 5W （Including Sensor unit） |  | POE standard arrangement
Dimensions | 70(W) × 120(H) × 150(D) |  | 
Weight | Approx. 0.9Kg (2.0 lbs) Including sensor unit |  | 
Mounting | Wall-mounting base plate by 3 screws |  | 
Sample Tubing | 4 x 6 mm PTFE tubing recommended. Ppfalf union fittings provided as standard accessories. |  | 
Cable | Cable type varies depending on communication method （Cable gland optional） |  | 
Notes: 12
12

## Slide 13: Specifications
Sensor Unit |  |  |  |  | 
Model | ESU | SGU | SSU | OSU | NCU
Detection Principle | Electrochemical Cell | Semiconductor | Pyrolysis-Particle | Galvanic Cell | New Ceramic
Identification Marking |  |  |  |  | 
Self-diagnosis Function | ・Sensor trouble ・System failure |  |  |  | 
Date-logger Function | ・Event history / Alarm history ・Calibration history ・Alarm trend （60 sec Before / After 1st alarm） |  |  |  | 
Notes: 13
13

## Slide 14: Specifications
Pyrolyzer Unit | 
Model | PLU-70
Application | Measuring NF3 / TEOS gases
Usage | Used by connecting to “GD-70D”(Main unit)
Display | ・LED (Green color)
 | - Normal：light-on / Warming-up ：flashing 1sec interval / Trouble： flashing 0.2 sec interval
Self-Diagnosis Function | ・Pyrolyzer unit trouble
 | ・Fan trouble
 | ・System trouble
Operational Temp. & Humidity | ・0-40℃ （Without rapid change） ・30-70% RH （non-condensing）
Operational Settings | Operational setting are user adjustable by using the front panel (Main unit).
Power Supply | DC 24V±10% Approx.36W (MAX)
Dimension | 70(W) × 120(H) × 150(D) (2.8”W x 4.7”H x 5.9”D)
Weight | Approx. 1.2Kg (2.6 lbs)
Mount-Type | Wall-mounting base plate by 3 screws
Sampling Tubing | 4 x 6 mm PTFE tubing recommended. PP half union fittings provided as standard accessories.
Cable | Cable type varies depending on communication method （Cable gland optional）
Notes: 14
14

## Slide 15: Power over Ethernet
PoE（Power over Ethernet） is a technology that can supply electricity to the detector through LAN cable, （IEEE802.3af）
With this PoE technology, data communication and electricity are available through one piece of LAN cable. Therefore, Detector (GD-70D) can be set without external power source, such as AC adaptor.
GD-70D requires less wiring and cost than units with additional power source（※CAT.5 LAN cable is necessary for PoE connection.
Gas detector(GD-70D) will become Ethernet terminal unit!
• Cable cut-down
（Line for electricity not required）
• Labor cost reduced
（Only RJ-45 connection required）
• Material cost reduced
（Cable for construction reduced）
Conventional connection
PoE connection
Supplying “data” and “electricity” through one LAN-cable.
AC100V / DC24V
Cable for Power-supply
LAN cable (Data only)
LAN cable (CAT.5)
Ethernet
Ethernet
Poe
Poe Hub
Notes: 15
15

## Slide 16: GD-70D Fresh Air
Press and hold “mode” for 3 seconds.
Press “SET”
Press “SET” to perform the zero adjustment
The menu displays “ZERO OK” if pass, or “ZERO NG” for fail
Notes: 16

## Slide 17: GD-70D Calibration
Press and hold “mode” for 5 seconds.
Press mode until 2-2.SPAN is displayed
Press the Set to select SPAN
Notes: 17

## Slide 18: GD-70D Calibration
Press and hold “mode” for 5 seconds.
Press the down arrow to select 1-8.M MODE
Press the Set to select 1-8.M MODE
Notes: 18

## Slide 19: GD-70D Calibration
Then press the SET key again for three seconds.
2.0 GAS TEST Perform a test with the gas. Similar to the detection mode, the reading changes and the alarm lamp lights up after the gas is introduced, but the contact is not activated.
Notes: 19

## Slide 20: GD-70D Calibration
2-1. ZERO Perform the zero adjustment.
2.2 GAS TEST Perform a test with the gas. Similar to the detection mode, the reading changes and the alarm lamp lights up after the gas is introduced, but the contact is not activated.
Press Set to begin Span Adjustment
Notes: 20

## Slide 21: GD-70D Calibration
Allow gas to flow, and then press the SET key when the reading is stabilized
Make Span Adjustment Because the reading will be fixed, select the introduced gas concentration by pressing the ▲ or ▼ key. Select the number, and then press the SET key. (20.9 vol% in the oxygen deficiency alarm specification)
Notes: 21

## Slide 22: GD-70D Calibration
Span adjustment completed
To record the adjustment result, press the SET key. (Press the MODE key to cancel this menu.)
The menu automatically returns to 2-2.SPAN
Notes: 22