## Slide 1: 03 Series Service PowerPoint

## Slide 2: 03 Series Components

## Slide 3: 03 Series Features
4 Gases to choose from: LEL, CO, H2S, or O2
Operates 3,000/2,000 hrs. on alkaline/Ni-MH (GP-03 35/30 hrs. on alkaline/Ni-MH)
Pocket size 2.1”(W) x 2.6”(H) x 0.9”(D)
Light weight 2.8 ounces
Audible / visual / vibration alarms
Automatic backlight during alarm
Peak Hold, STEL & TWA and Low battery alarm
Impact and water resistant (IP-67)
Intrinsically safe, ATEX, CSA pending
Replaceable boot
Datalogging
Bump/Calibration reminder

## Slide 4: Applications
Hazardous Material
Mining
Refineries / Petrochemical
Water / Wastewater Treatment
Pharmaceuticals
Utilities
Chemical Plants
Fire Service

## Slide 5: 03 Series Control Buttons

## Slide 6: Start-up Procedure
Press and hold the POWER MODE button
  Release when the instrument beeps
    The initial startup screen depends on how the following parameters are set (using the 03 Series User Setup Program):
      Bump Fail Behavior
      Cal. Limit Display
      Cal. Limit Check
      Auto Zero Adjustment

## Slide 7: Start-up Procedure Cont.
The display then indicates the following items for about a second each:
Full scale value
Warning set point (low gas alarm)
Alarm set point (high gas alarm)
STEL alarm set point (CO and H2S only)
TWA alarm set point (CO and H2S only)
The 03 Series is now operating in Measuring Mode and monitoring for gas

## Slide 8: Fresh Air Adjustment
Verify that the instrument is in a fresh air environment
Press and hold the AIR button
Release when the display says “ADJ”
When the fresh air adjustment is finished the instrument will return to the Setup Mode

## Slide 9: Adjusting the Buzzer Volume
In Normal Measuring mode press and hold the POWER MODE and AIR buttons simultaneously
  To avoid entering Display Mode, press the AIR button slightly sooner than pressing the POWER MODE button
Release both buttons when the instrument starts to simulate an alarm
Use the AIR button to select LO or HI
Press and release the POWER MODE button to save the setting and return to the normal measuring mode
  If LO is selected it will appear at the bottom of the normal operation screen

## Slide 10: Display Mode
Press the POWER MODE button in Normal mode
  Peak
  STEL (Only HS-03 and CO-03)
  TWA (Only HS-03 and CO-03)
  Full Scale
  Date & Time

## Slide 11: Setup Mode
Allows access to:
  Date & Time (DATE)
  Fresh Air Adjustment (AIR)
  Automatic Span Adjustment (A--CAL)
  Easy Span Adjustment (E--CAL)
  Manual Span Adjustment (M--CAL)
  Alarm Points (ALM--P)
  Bump Test Limit Check Setting (BP--CHK)
  Password (PASS--W)

## Slide 12: Using the Setup Mode
Make sure the 03 Series is OFF
Press and hold the AIR button, then press and hold the POWER MODE button.
Wait until you hear two beeps, then Release both buttons
  If the PASS--W menu item is set to On (located in the setup Mode), use the AIR button to scroll to the correct number
  Press the POWER MODE to move to the next digit, and when all the digits are entered press the POWER MODE to enter the setup mode

## Slide 13: Setting the Date & Time
Press the POWER MODE button when the DATE screen is displayed
Use the AIR button to get to the desired year
Press the POWER MODE button to save the setting
Continue the previous two steps to change the month, day, hour and minute setting

## Slide 14: Performing an Automatic Span Adjustment
Press the POWER MODE button when the A--CAL screen is displayed
The calibration gas concentration will be displayed
  If the concentration does not match, Press and hold the AIR button, then press and hold the POWER MODE button, release both when you here a beep.
  Use the AIR button to adjust the value
  Press the POWER MODE button once the correct value is displayed
  The unit will return to the previous screen
Press the POWER MODE button
  LCD will display the current gas readings and “A--CAL.” will flash

## Slide 15: Performing an Automatic Span Adjustment
Connect the regulator, cylinder, calibration tubing and calibration cup together
Connect the calibration cup to the instrument
Turn on the regulator
Allow the gas to flow for 2 minutes
Press the POWER MODE button
If the span adjustment is successful the unit will show PASS, if the span adjustment is not successful, the unit will show FAIL
Turn off regulator and disconnect the calibration cup

## Slide 16: Performing a Easy Span Adjustment
Press the POWER MODE button when the E--CAL screen is displayed
The calibration gas concentration will be displayed
  If the concentration does not match, Press and hold the AIR button, then press and hold the POWER MODE button, release both when you here a beep.
  Use the AIR button to adjust the value
  Press the POWER MODE button once the correct value is displayed
  The unit will return to the previous screen
Press the POWER MODE button
  LCD will display the current gas readings and “E--CAL.” will flash

## Slide 17: Performing a Easy Span Adjustment
Connect the regulator, cylinder, calibration tubing and calibration cup together
Connect the calibration cup to the instrument
Turn on the regulator
Once the gas reading reaches 10% on the full scale value, E--CAL will begin to flash and a time countdown will begin

## Slide 18: Performing a Easy Span Adjustment
At the end of the countdown, the 03 Series will attempt to make a span adjustment (zero adjustment for OX-03)
  If the span adjustment (zero adjustment for OX-03) is successful, the LCD will show the following screen before returning to the E--CAL screen
  If the span adjustment (zero adjustment for OX-03) fails, the LCD will show Fail
Turn off regulator and disconnect the calibration cup
Press and release the POWER MODE button to clear the failure indication and return to the E--CAL screen

## Slide 19: Performing a Manual Span Adjustment
Press the POWER MODE button when the M--CAL screen is displayed
Press the POWER MODE button
  LCD will display the current gas readings and “M--CAL.” will flash
Connect the regulator, cylinder, calibration tubing and calibration cup together
Connect the calibration cup to the instrument
Turn on the regulator
Allow the gas to flow for 2 minutes
Adjust the display gas reading using the AIR button to match the concentration of the calibration cylinder
Press the POWER MODE button
After the span adjustment is made, the instrument will return to the M--CAL screen
Turn off regulator and disconnect the calibration cup

## Slide 20: Performing a Bump Test
Press the POWER MODE button when the BUMP screen is displayed
Connect the regulator, cylinder, calibration tubing together
Connect the calibration cup to the instrument
Connect the other end of the calibration tubing to the calibration cup
Press the POWER MODE button
  LCD will alternate between BUMP and APPLY
Turn on the regulator
When the bump test countdown reaches 0, the unit will let you know whether it passed or failed

## Slide 21: Performing a Bump Test
Press AIR to view the bump test gas reading
To return to the BUMP screen press the POWER MODE button
If Calibration after Bump Test Fails is set to ON a calibration is immediately and automatically started upon a bump test failure
The results are displayed after countdown finishes

## Slide 22: Maintenance Mode
Allows access to:
  Easy Calibration (E--CAL)
  Manual Span Adjustment (M--CAL)
  ROM
    firmware version that is loaded in the instrument and the firmware checksum
  Normal Mode (START)
  Date & Time (DATE)
  Fresh Air Adjustment (AIR)
With the Unit off press and hold the AIR and Power Buttons for one beep

## Slide 23: Factory Mode
Allows access to:
  Changing the Combustible Gas in the GP-03 (COMB)
  Alarm Points (ALM--P)
  Latching (LATCH)
    Latching/non-latching
  ROM
  A/D Value (A/D)
  Display Alarm Points (D--AL--P)
    Full scale/Alarm point (alarm test)
  Default (DEF.)
  Normal Mode (START)
With the Unit off press and hold the AIR and Power Buttons for three beeps Password is 1994, use the AIR button to get to the correct digit POWER MODE button to get to the next digit

## Slide 24: Replacing the Filters
Turn the 03 Series is off
Remove the Rubber Boot
Unscrew the four screws that attach the rear case to the front case
Turn the instrument right side up (The screws are not captive and will fall out)
Remove the front case from the rear case
Carefully remove the sensor gasket (If you have a CO-03, the charcoal filter will be installed in the sensor gasket)
Carefully remove the sensor membrane
If you have a CO-03, install a charcoal filter in to the recess in the sensor gasket.
Then insert a new sensor membrane in the sensor membrane recess
Reinstall the sensor gasket with the flat side down.
Turn the front case right side up and carefully secure it to the rear case
Reinstall the rubber boot if it is being used

## Slide 25: Disassembly of Unit
Rotate the captive battery cover screw counter clockwise to remove the battery cover

## Slide 26: Disassembly of Unit
Carefully unscrew the four screws that attach the rear case to the front case

## Slide 27: Disassembly of Unit
Remove the front case from the rear case
Caution: Circuit boards may be damaged by ESD! Work must be performed in ESD controlled environment.

## Slide 28: Disassembly of Unit
Remove the Sensor
Caution: Circuit boards may be damaged by ESD! Work must be performed in ESD controlled environment.

## Slide 29: Disassembly of Unit
Gently pry up the four corners of the LCD screen (no screws are holding it in)
Caution: Circuit boards may be damaged by ESD! Work must be performed in ESD controlled environment.

## Slide 30: Disassembly of Unit
Caution: Circuit boards may be damaged by ESD! Work must be performed in ESD controlled environment.

## Slide 31: Disassembly of Unit
Slide Latch up to release ribbon cable pins
Caution: Circuit boards may be damaged by ESD! Work must be performed in ESD controlled environment.

## Slide 32: Disassembly of Unit
Remove Vibration Motor

## Slide 33: Disassembly of Unit

## Slide 34: Disassembly of Unit

## Slide 35
Reassemble in reverse order

## Slide 36: Questions?