## Slide 1
GX-3R User Training

## Slide 2
Turn on the GX-3R by pressing and holding the POWER MODE button until you hear one beep
Turn on

## Slide 3
Press the POWER/MODE button to perform calibration.
Cal Reminder
Acknowledge and use
Perform a calibration
Press AIR button to continue without performing a calibration
If your instrument is past its 30 day calibration limit, you will see a CAL-LMT reminder at start-up

## Slide 4
After warm-up, the normal screen will appear.
Normal Mode

## Slide 5
Fresh Air
To perform a fresh air adjustment, make sure you’re in a fresh air environment.

## Slide 6
Press and hold the AIR button
Fresh Air

## Slide 7
When prompted, release the AIR button
Fresh Air
The GX-3R will beep once and return to normal mode after the fresh air adjust is complete

## Slide 8
To reset the instrument, go to a safe location. Once the readings return to fresh air values, then press the POWER/MODE button.
Alarm
In the event of an alarm, the GX-3R’s visual, audible and vibration alarms will alert you.
The readout will identify which sensor is in alarm.
The visual, audible, and vibration alarms will stop.

## Slide 9: Display Mode
To enter display mode, press and release the Power Mode button

## Slide 10
You will know you are in display mode if you see a display mode menu title on the bottom of your screen
Display mode vs. Normal Operation
PEAK
STEL
TWA
HC GAS LIST
USER ID
STN ID
CAL DATA
BUMP DATA
TEMPERATURE
ALARM-PT
To return to normal operation, scroll all the way through the display menu using the AIR button, or wait 20 seconds and the GX-3R will automatically return to normal mode

## Slide 11
Charging the instrument
Plug the charger into the instrument. A full charge takes 3 hours, and will power the instrument up to 25 hours.
A 30 minute charge will provide 8 hours of operation time.
Once the GX-3R is charged, the light will go from orange to green
Multi-unit AC Adapter is available (#49-0134)

## Slide 12
Downloading Data
Three methods to downloading data from the GX-3R
IrDA cable to PC
SDM-3R to PC
SDM-3R to flash drive to PC

## Slide 13
IrDA cable to PC
Launch the GX-3R Datalogging program
Align the infrared ports
Turn on GX-3R
IrDA Cable
For More detailed instructions: www.gasdetecting.com/files/GX3R_IrDA_to_PC.pdf

## Slide 14
SDM-3R to PC
Launch the SDM-3R PC Controller
Place GX-3R in the SDM-3R docking station
Turn on SDM-3R
SDM-3R to PC
For More detailed instructions: www.gasdetecting.com/files/SDM3R_to_flashdrive_to_PC.pdf

## Slide 15
SDM-3R to Flash drive to PC
Insert flash drive into the SDM-3R and press COPY
Launch the SDM-3R PC Controller
Insert flash drive into the computer’s USB slot
Flash drive
For More detailed instructions: www.gasdetecting.com/files/SDM3R_to_PC.pdf

## Slide 16
RP-3R Attachable Pump

## Slide 17
RP-3R Pump
Place the GX-3R right bottom first (as illustrated above) and then seat the left side until the clip latches.

## Slide 18
RP-3R Pump
The attachable probe has a dust filter to ensure the quality of your samples
The inlet filter housing contains a hydrophobic filter

## Slide 19
RP-3R Pump
Low flow alarm
If the sample line is blocked, the amber flow LED will flash and the buzzer with intermittently sound

## Slide 20
RP-3R Pump
Other Alarm Types

## Slide 21
Turn off the GX-3R by pressing and holding the POWER MODE button for approximately 5 seconds
Turn off

## Slide 22
SDM-3R Calibration Station

## Slide 23
SDM-3R: Gas Ports
Air port:
Dust Filter
Gas 1:
4 gas mix
Gas 2:
Not used
Gas 3
Not used

## Slide 24
SDM-3R: USB port
USB port

## Slide 25
SDM-3R: Rear ports
AC power
Ethernet
PC Connection

## Slide 26
Connect Cylinder to SDM-3R
Connect a 4 gas mix cylinder into “Gas 1 “ port.
Make sure to use a demand flow regulator

## Slide 27
SDM-3R: GX-3R Placement
Plug in power adaptor and turn on the SDM-3R
Make sure the GX-3R is powered off before installing into the SDM-3R
When placing the GX-3R into the docking station, remember to have the yellow tab engaged upwards

## Slide 28
Turn on
If an automatic test is set up to occur, it will start as soon as the GX-3R is connected to the SDM-3R
Close the red door completely, the GX-3R will turn on automatically
When the GX-3R is connected to the SDM-3R, the GX-3R’s date and time are automatically updated to the SDM-3R’s current time and date

## Slide 29
Bump Test
Press the BUMP button to start bump test
Note: Verify that the appropriate calibration gas cylinder is connected to the gas 1 fitting on the left side of the SDM-3R

## Slide 30
Bump Test
During the bump test, the GX-3R’s screen will display the current gas readings

## Slide 31
Bump Test
Once a bump test is complete, the LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor.
The “BUMP” LED light on the SDM-3R will be green when a bump test passes.
The GX-3R will shut off automatically after successful bump test.

## Slide 32
Bump Test
If the bump test fails, ”BUMP” LED light is red. The SDM-3R will then automatically perform a calibration
The LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor for both bump test and calibration
If calibration is successful, the GX-3R will automatically power off

## Slide 33
Calibration
Press the CAL button until you hear the pump to start calibration
Make sure calibration gas concentrations match values shown on display

## Slide 34
Calibration
During calibration, the display shows current gas readings.
Calibration will take approximately 90 seconds.

## Slide 35
Calibration
Once calibration is complete, the LCD screen will alternate between the gas concentration and the letter “P” for pass or an “F” for fail for each sensor
The “CAL” LED light on the SDM-3R will be green when calibration is passed
“CAL” LED light is red if there if a failure. The GX-3R will power off after a successful calibration

## Slide 36
Setting Date/time
Insert GX-3R into the SDM-3R docking station
Press POWER

## Slide 37
Setting Date/time
Press EDIT/ENTER to enter settings menu
Press BUMP to scroll to date and press EDIT/ENTER

## Slide 38
Setting Date/time
Use BUMP and CAL to scroll to the date or time, and press EDIT/ENTER
Use BUMP and CAL to change the value, and press EDIT/ENTER to accept changes.

## Slide 39
Removing from SDM-3R
Release the tab on the red door
Remove the GX-3R

## Slide 40: Training Resources
SDM-3R Cal Station Poster
Ships with each SDM-3R

## Slide 41: Training Resources
YouTube GX-3R Play List
  https://www.youtube.com/playlist

## Slide 42: Training Resources
GX-3R webpage

## Slide 43: Training Resources
RKI Regional Representatives hosted webinars

## Slide 44
Questions?