## Slide 1
GX-3R Pro User Training

## Slide 2
Turn on the GX-3R Pro by pressing and holding the POWER MODE button until you hear one beep
Turn on

## Slide 3
Press the POWER/MODE button to perform calibration.
Cal Reminder
Acknowledge and use
Perform a calibration
Press AIR button to continue without performing a calibration
If your calibration is due, you will see a Cal Reminder upon warm-up

## Slide 4
After warm-up, the normal screen will appear.
Normal Mode

## Slide 5
Fresh Air
To perform a fresh air adjustment, make sure you’re in a fresh air environment.

## Slide 6
Press and hold the AIR button
Fresh Air

## Slide 7
When prompted, release the AIR button
Fresh Air
The GX-3R Pro will beep once and return to normal mode after the fresh air adjust is complete

## Slide 8
To reset the instrument, go to a safe location. Once the readings return to fresh air values, then press the POWER/MODE button.
Alarm
In the event of an alarm, the GX-3R Pro’s visual, audible and vibration alarms will alert you.
The readout will identify which sensors are in alarm.
The visual, audible, and vibration alarms will stop.

## Slide 9: Display Mode
To enter display mode, press and release the Power Mode button

## Slide 10
You will know you are in display mode if you see a display mode menu title on the bottom of your screen
Display mode vs. Normal Operation
PEAK
STEL
TWA
HC GAS LIST
USER ID
STN ID
CAL DATA
BUMP DATA
TEMPERATURE
ALARM-PT
To return to normal operation, scroll all the way through the display menu using the POWER/MODE button, or wait 20 seconds and the GX-3R Pro will automatically return to normal mode

## Slide 11
Charging the instrument
Plug the charger into the instrument. A full charge takes 3 hours, and will power the instrument up to 25 hours.
A 30 minute charge will provide 8 hours of operation time.
Once the GX-3R Pro is charged, the light will go from orange to green
Multi-unit AC Adapter is available (#49-0134)

## Slide 12
Turn off the GX-3R Pro by pressing and holding the POWER MODE button
Turn off

## Slide 13
RK Link App can be configured while connected or not connected to a GX-3R Pro instrument
RK Link APP
Instrument Connected
Instrument Not Connected

## Slide 14
Select instrument available to pair
Then select the Pair button
Pair any GX-3R Pro while the instrument’s Bluetooth is powered on.
Pair to any GX-3R Pro
Choose Pair Button
Instrument Connects

## Slide 15
RK Link App can send email and / or text alarm notifications.
Email/Text Notifications
Email Notifications
Text Notifications

## Slide 16
Register with the RK Link server in order to begin text or email notifications.
Email/Text Notifications
Choose Application Settings
Choose Notification Settings
Choose Main Menu

## Slide 17
Do not enter any Gmail info
(for international use)
Complete form with no dashes on phone #
Use Back to return to Application Settings menu
Choose Application Server Settings
From the Application Server Settings window, you will register with the RK Link server and receive a one-time User ID.
Email/Text Notifications
Now ready to enter email / texts contacts

## Slide 18
RK Link can direct different alarm types to separate notification contact sets. 5 total alarm triggers available to chose.
Set Types of Alarm Triggers
Create & set alarm types to notify
Choose Email notification settings
Touch the plus button to load email and/or text numbers

## Slide 19
Enter mobile number with 1 and no dashes.
Touch the New Registration button to accept the entry.
RK Link can send both email and or text messages.
Entering text / email
Touch New Registration button to clear keyboard.
Enter email then touch New Registration button.

## Slide 20
Touch New registration button to clear keyboard
Select the entered contacts to receive text or email alarm notifications.
Selecting text / email
Touch the OK button to accept set contacts
Touch the entered email or cell number to select contact

## Slide 21
Live gas readings
You’re now configured for email / text alarm notifications.
Now Ready for Notifications
Use the back button to return to gas readings

## Slide 22
Unscrew the captive screws on the GX-3R Pro.
Sensor/Filter replacement

## Slide 23
Sensor/Filter replacement
O2 Sensor:
ESR-X13P
LEL Sensor:
NCR-6309
H2S/CO:
ESR-A1DP
Sensor Dust Filter:
33-0182-10 (10 pk)
Scrubber Disk Filter for H2S/CO:
33-7130 (5 pk)
Scrubber Disk Filter for LEL:
33-7131 (5 pk)
Toxic Sensor:
SO2: ESR-A13D
Scrubber Disk Filter for SO2:

## Slide 24
RP-3R Attachable Pump

## Slide 25
RP-3R Pump
Place the GX-3R Pro’s right bottom first (as illustrated above) and then seat the left side until the clip latches.

## Slide 26
RP-3R Pump
The attachable probe has a dust filter to ensure the quality of your samples
The inlet filter housing contains a hydrophobic filter

## Slide 27
SDM-3R Calibration Station

## Slide 28
SDM-3R: Gas Ports
Air port:
Dust Filter
Gas 1:
4 gas mix
Gas 2:
SO2
Gas 3
Not used

## Slide 29
SDM-3R: USB port
USB port

## Slide 30
SDM-3R: Rear ports
AC power
Ethernet
PC Connection

## Slide 31
Connect Cylinder to SDM-3R
Connect a 4 gas mix cylinder into “Gas 1 “ port.
Make sure to use a demand flow regulator

## Slide 32
SDM-3R: GX-3R Placement
Plug in power adaptor and turn on the SDM-3R
Make sure the GX-3R Pro is powered off before installing into the SDM-3R
When placing the GX-3R Pro into the docking station, remember to have the yellow tab engaged downwards

## Slide 33
Turn on
If an automatic test is set up to occur, it will start as soon as the GX-3R Pro is connected to the SDM-3R
Close the red door completely, the GX-3R Pro will turn on automatically
When the GX-3R Pro is connected to the SDM-3R, the GX-3R Pro’s date and time are automatically updated to the SDM-3R’s current time and date

## Slide 34
Bump Test
Press the BUMP button to start bump test
Note: Verify that the appropriate calibration gas cylinder is connected to the gas 1 fitting on the left side of the SDM-3R

## Slide 35
Bump Test
During the bump test, the GX-3R Pro’s screen will display the current gas readings

## Slide 36
Bump Test
Once a bump test is complete, the LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor.
The “BUMP” LED light on the SDM-3R will be green when a bump test passes.
The GX-3R Pro will shut off automatically after successful bump test.

## Slide 37
Bump Test
If the bump test fails, ”BUMP” LED light is red. The SDM-3R will then automatically perform a calibration
The LCD screen will alternate between the gas concentration and pass/fail status (shown as P or F on screen) for each sensor for both bump test and calibration
If calibration is successful, the GX-3R will automatically power off

## Slide 38
Calibration
Press the CAL button until you hear the pump to start calibration
Make sure calibration gas concentrations match values shown on display

## Slide 39
Calibration
During calibration, the display shows current gas readings.
Calibration will take approximately 90 seconds.

## Slide 40
Calibration
Once calibration is complete, the LCD screen will alternate between the gas concentration and the letter “P” for pass or an “F” for fail for each sensor
The “CAL” LED light on the SDM-3R will be green when calibration is passed
“CAL” LED light is red if there if a failure. The GX-3R will power off after a successful calibration

## Slide 41
Removing from SDM-3R
Release the tab on the red door
Remove the GX-3R

## Slide 42
Data Transfer
Insert a USB flash drive with unit powered on.
To download the calibration and bump test data, press and hold COPY until the COPY LED is solid red, then release.
To download stored data from the gas monitor, keep pressing the COPY button until LED is flashing orange.
The USB flash drive can be removed when the COPY LED flashes green

## Slide 43
Questions?