## Slide 1
GX-6000
6 gas, sample draw with PID & Super Toxics
Removing Main PCB

## Slide 2
Remove battery and sensor cover from instrument.
Remove rubber boot and scrubbers.
Remove the sensors including sensors that are installed into the Smart 1 and Smart 2 locations.
Removing the Main PCB

## Slide 3
Unscrew the four Phillips screws securing the top of the detector.
Remove the top exposing the pump.
Grasp the pump and remove it from the unit.
Removing the Main PCB

## Slide 4
Remove screw as shown.
Removing the Main PCB

## Slide 5
Pull sensor chamber out and remove ribbon connector from main circuit board.
Removing the Main PCB

## Slide 6
Lift second chamber as shown
Removing the Main PCB

## Slide 7
Chamber and ribbon cable have been removed.
Remove rubber joint for pressure sensor. P/N 07-6039
Removing the Main PCB

## Slide 8
All parts removed.
Removing the Main PCB

## Slide 9
Gently slide circuit board towards the top of the instrument to remove.
Take care not to damage battery spring contacts.
Removing the Main PCB

## Slide 10
Main PCB with LCD shown.
Buzzer facing up.
Removing the Main PCB

## Slide 11
LCD connector has a lock that needs to be released before ribbon cable can be removed.
Memory battery shown can be easily replaced.
P/N 49-1050RK
Removing the Main PCB

## Slide 12
LCD P/N 51-1139
Clean LCD before installing to remove any fingerprints.
Removing the Main PCB

## Slide 13
Main PCB, Component side.
P/N 57-1286
Removing the Main PCB

## Slide 14
Main PCB, sensor side shown.
Differential Pressure switch.
Removing the Main PCB

## Slide 15: Questions?