Flow adjustment for GX-8000
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1) Go into the “FACTORY MODE”
Hold "AIR" + "RESET" and Push "POWER ENTER".  Input the password “1994” by using "AIR" + "RESET" buttons and press “ENTER”.

2) Go into the mode “FLOW-ADJ” with “▲” and ”▼” buttons and press “ENTER”.
3) Adjust the value “1500 +/- 100” by “▲” and “▼” buttons and press “ENTER”.  (Standardization for the output of the pressure sensor in the pump OFF condition)
4) Check if the pump ON value is between 1700 and 2000.

5) Adjust the flow rate to 350 +/- 30ml/min by needle valve and press “ENTER” to set.
6) The message “END” is displayed and “FLOW ADJ” mode is automatically returned.
7) Select “START” by “▲” and ”▼” buttons and enter measuring mode.

* Please make sure that FAIL FLOW error is not appeared on the display.
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