## Slide 1
GX-Force
Service PowerPoint

## Slide 3
Agenda
Control Buttons
Start-up procedure
Fresh air adjust
Display Mode
Calibration
Filter & sensor replacement

## Slide 4
Button | Function(s)
AIR | Turns on LCD backlight Reset alarm button for latching alarms Enters User Mode, Maintenance Mode, and Gas Select mode when used with POWER MODE Activates the demand zero for fresh air reading Changes the value of a parameter available for adjustment Scrolls through parameter options
Power MODE | Turns the GX-Force on and off Turns on LCD backlight Enters and scrolls through Display Mode Save button Reset alarm button for latching alarms Enters User Mode, Maintenance Mode, and Gas Select mode when used with AIR
GX-Force
Control Buttons

## Slide 5
GX-Force
Start-up Procedure
Press and hold POWER MODE button. Release the button when you hear a beep.

## Slide 6
GX-Force
Start-up Procedure
Start-up sequence will commence
Press the AIR button to bypass the CAL.RMDR and continue using the GX-Force
If CAL.RMDR is set to ON (factory setting) and a calibration is due, the screen that appears next depends on the CAL.EXPD setting in User Mode.
Press the POWER/MODE button to calibrate the instrument in User Mode
Date and Time

## Slide 7
GX-Force
Start-up Procedure
Full Scale
Target Gases
Battery Voltage

## Slide 8
GX-Force
Start-up Procedure
High Alarm
Alarm
Warning

## Slide 9
GX-Force
Start-up Procedure
Normal Operation Screen
TWA
STEL

## Slide 10
GX-Force
Fresh Air Adjust
Verify that the instrument is in
a fresh air environment
Press and hold the AIR button
Keep holding the AIR button until prompted to release it

## Slide 11
GX-Force
Display Mode
From the Normal Operation Screen ( Measuring Mode ) Press POWER/MODE to cycle through Display Mode’s menu items
Press AIR to turn on/off the light
Flashlight
STEL
Peak

## Slide 12
GX-Force
Display Mode
TWA
HC Gas List
Press AIR to change HC Gas List
Press POWER/MODE to continue to the next display menu

## Slide 13
GX-Force
Display Mode
Calibration Data
Press AIR to display cal. data
Press POWER/MODE to continue to the next display menu

## Slide 14
GX-Force
Display Mode
Snap Log Data
Press AIR to display snap log data
Press POWER/MODE to continue to the next display menu

## Slide 15
GX-Force
Display Mode
Date / Time / Temp
Alarm Points
Press AIR to display alarm points
Press POWER/MODE to continue to the next display menu
Normal Operation Screen

## Slide 16
GX-Force
Calibration (GAS CAL)
Calibration Cylinder
Demand Flow Regulator
Tubing

## Slide 17
GX-Force
Calibration (GAS CAL)
With the instrument off, Press AIR and then POWER/MODE to enter User Mode
BUMP is the first menu to appear
Press AIR to scroll to GAS CAL and then press POWER/MODE

## Slide 18
GX-Force
Calibration (GAS CAL)
AIR CAL appears on the screen. Press POWER/MODE
Press and hold AIR until prompted to release it

## Slide 19
GX-Force
Calibration (GAS CAL)
Press AIR to scroll to AUTOCAL
Press POWER/MODE to enter AUTOCAL
Connect the demand flow regulator to the calibration cylinder, attach the hose to the demand flow regulator and the instrument

## Slide 20
GX-Force
Calibration (GAS CAL)
The screen will show the calibration values, which should match the calibration cylinder. Press POWER/ MODE, and let the gas flow for 1 minute
After the reading stabilizes, press POWER/MODE. The result will show PASS. The instrument will automatically start up to Normal Operation
If the result shows FAIL, press POWER /MODE. The instrument will prompt you to repeat the calibration process

## Slide 21
GX-Force
Leak Check Mode
Press AIR to scroll to the Leak Check Mode item
With the Leak Check Mode selection displayed, press POWER MODE to begin the Leak Check Mode startup sequence
Press and briefly hold down POWER MODE. Release the button when you hear a beep. The Mode Select Screen displays
With the Leak Check Mode displayed, you are ready to use the GX-Force

## Slide 22
GX-Force
Leak Check Mode
From the Leak Check Operation Screen Press POWER/MODE to cycle through Display Mode’s menu items
Flashlight
Peak
Snap Log Data

## Slide 23
GX-Force
Adjusting Leak Check Settings
Navigate to the RKI website
Select “Product Resources”, then click on “Software”
Download the GX-Force Setup Program to adjust Leak Check Mode parameters

## Slide 24
GX-Force
Adjusting Leak Check Settings
On the Software page, click “Portable Monitors”, then click on the “GX-Force Setup Program” link to begin downloading the installation file

## Slide 25
GX-Force
Adjusting Leak Check Settings
Open the GX-Force Setup program setup.exe file and follow the on-screen installation guide
Connect the GX-Force to the computer using a Micro USB Type C cable
Open the GX-Force Setup Program

## Slide 26
GX-Force
Adjusting Leak Check Settings
Enter 1939 into the password input field. The following interface will appear.

## Slide 27
GX-Force
Adjusting Leak Check Settings
Press POWER MODE to turn on the instrument. The unit will begin communicating with the computer

## Slide 28
GX-Force
Adjusting Leak Check Settings
Both the instrument screen and status area in the Setup Program will update accordingly

## Slide 29
GX-Force
Adjusting Leak Check Settings
Click on the “Instrument Information” button in the GX-Force Program

## Slide 30
GX-Force
Adjusting Leak Check Settings
Click on the “Detail Settings” button to open the Parameters window

## Slide 31
GX-Force
Adjusting Leak Check Settings
The parameter window displays the Maintenance Tab
To adjust the parameter, double click on the right column of the setting
NC (factory setting): Only the LEL channel displays in Leak Check Mode
NC+CO : The LEL and CO channels are displayed in Leak Check Mode
NC+O2 : The LEL and O2 channels are displayed in Leak Check Mode
NC+CO+O2 : The LEL, CO, and O2 channels are displayed in Leak Check Mode

## Slide 32
GX-Force
Adjusting Leak Check Settings
Press “Update” in the main Setup window to save all adjustments made to the instrument

## Slide 33
GX-Force
Adjusting Leak Check Settings
Click the “Power Off” button to power off the instrument

## Slide 34
Filter & Sensor Replacement
Take the GX-Force to a non-static environment and turn it off if it is on.
GX-Force
Unscrew the four captive screws

## Slide 35
Filter & Sensor Replacement
GX-Force
Lift up the Rear Case Assembly and set it aside
Part # 21-1982

## Slide 36
Filter & Sensor Replacement
GX-Force
Lift up the Filter Gasket, and replace
if needed. Part # 07-6064

## Slide 37
Filter & Sensor Replacement
GX-Force
Lift up the Sensor Filters, and replace if needed
Filter Part #’s
33-7130:
33-7131: 33-7132: 33-7133:
Scrubber disk filter for dual CO/H2S Scrubber disk filter for LEL
Scrubber disk filter for CO and CO(-H2) Humidity filter disk for H2S

## Slide 38
Filter & Sensor Replacement
GX-Force
Lift up the Sensor Retainer, and replace
if needed. Part # 21-1983

## Slide 39
Filter & Sensor Replacement
GX-Force
Lift up the Sensors, and replace if needed
Sensor Part #’s
ESR-A13i-H2S: ESR-A1DP-COHS: ESR-A13P-CO: ESR-A1CP-CO-H: ESR-X13P-OXY: NCR-6309:
H2S Sensor CO/ H2S Combo Sensor CO Sensor CO with H2 comp. O2 Sensor LEL Sensor

## Slide 40
Filter & Sensor Replacement
GX-Force
Remove the Rubber Nozzle and replace if needed. Part # 17-1001RK
Remove the Filter Holder (twist it counter clockwise) and replace if needed. Part # 21-1833RK

## Slide 41
Filter & Sensor Replacement
GX-Force
Remove the Teflon Filter Disk and Wire Mesh Disk Filter and replace if needed. Part # 33-0159RK / 33-1112RK
Remove the Gasket and replace if needed. Part # 07-2005RK